Clinical Evaluation Plan for

Axium Neurostimulator System

Approval:

| Role | Approver | Signature/Date (ddMMMyyyy) |
| --- | --- | --- |
| Global Regulatory Operations-Clinical Evaluation (GRO-CE) /Author | Amy Liu | See electronic signature in Windchill |
| Global Regulatory Operations-Clinical Evaluation (GRO-CE) Medical Writer | Mary Fautsch | See electronic signature in Windchill |
| GRO-CE PM/Director | Sangeeta Mantripragada | See electronic signature in Windchill |
| R&D/Clinical Engineering | Tj Baggett | See electronic signature in Windchill |
| Regulatory Affairs | Jordan Hanson | See electronic signature in Windchill |
| Medical Affairs | Allen Burton, MD | See electronic signature in Windchill |
| Clinical Affairs | Edward Karst | See electronic signature in Windchill |

# Revision History

| Revision | Description |
| --- | --- |
| A | Initial release |
| B | Annual update, Annual update to meet requirements of 87281, Rev C, including update of post market surveillance data, SOA and literature review to include relevant clinical data from articles published from November 1, 2017-December 15, 2018 |
| C | Annual update; Annual update to meet requirements of 87281, Rev D, including update of post market surveillance data, SOA and literature review to include relevant clinical data from articles published from December 15, 2018-December 31, 2019 |

# Table of Contents

1.0	Scope of Clinical Evaluation Plan (Stage 0)	5

1.1	Objective	5

1.2	Device(s) under Evaluation	6

1.3	Type of Assessment	6

1.4	Timeframe of Assessment	7

1.5	Newly Emerged Clinical Concerns	8

1.6	Summary of Prior Clinical Evaluation Activities	8

1.7	Relevant <Essential Requirements> OR <General Safety and Performance Requirements>	8

1.8	Applicable Regulatory References	9

2.0	Clinical Background and State of the Art	10

2.1	Literature Search Summary	12

2.1.1	Background	12

2.1.2	Objectives	12

2.1.3	Methods	12

2.1.4	Results	12

2.2	Literature Review	15

2.2.1	Medical Conditions	15

2.2.2	Technological Developments	15

2.2.3	Therapeutic, Diagnostic, or Treatment Options Summary	16

2.2.4	Similar Devices	17

2.2.5	State of the Art Landscape Analysis	17

2.2.6	Conclusion	20

3.0	Safety and Performance Measures/Objectives	20

4.0	Device Description	23

4.1	Detailed Device Description	23

4.2	Name and Address of Manufacturer	24

4.3	Principles of Operation	24

4.4	Development History	24

4.5	Regulatory History	24

4.6	Product Labeling	25

4.6.1	Intended Purpose/Use	25

4.6.2	Intended Target Groups	25

4.6.3	Indications	25

4.6.4	Contraindications	25

4.6.5	Warnings	25

4.6.6	Precautions	25

4.6.7	Potential Adverse Events	26

4.7	Intended User	26

4.8	Intended Clinical Benefits	26

4.9	Claims	26

4.10	Devices Used in Combination (if applicable)	26

4.11	Animal Tissues or Derivatives (if applicable)	26

5.0	Clinical Development Plan (Stage 0)	27

6.0	Equivalency Assessment (Stage 0)	27

6.1	Equivalent Device(s)	29

6.2	Equivalence Tables	29

7.0	Identify Pertinent Data for Clinical Evaluation (stage 1)	34

7.1	Data Sources	34

7.2	Non-Clinical Data	34

7.3	Clinical Investigations (or Justification for No Clinical Investigation)	35

7.4	Data Retrieved from Literature	35

7.4.1	Literature Search Protocol	35

7.5	Post-Market Surveillance Data	36

7.6	Post-Market Clinical Follow-Up Data	37

8.0	Appraisal of Pertinent Data for Clinical Evaluation (STAGE 2)	37

8.1.1	Appraisal Criteria	37

8.1.2	Analysis Criteria for Quality of Clinical Data	38

9.0	Analysis of Clinical Data (Benefit-Risk Analysis) (STAGE 3)	39

9.1	Safety and Performance	39

9.2	Risk Analysis (from Risk Management Documentation)	40

9.3	Benefit-Risk Assessment	40

9.4	Additional Data	41

10.0	Summary of the Clinical Evaluation (Stage 4)	42

11.0	Updating the Clinical Evaluation Plan	43

# Table of Tables

Table 1: Axium Products (Current AIMDD Certified)	6

Table 2: Axium Products (Rationalized from EU)	7

Table 3. Axium CER Revision History	9

Table 4: Regulatory Approval History (Current AIMDD Axium devices)	14

Table 5: Regulatory Approval History (Current AIMDD devices used with Proclaim DRG System)	15

Table 6: Regulatory History of Rationalized Axium Devices	15

Table 7 Countries Where Axium Neurostimulation System Approved	16

Table 8: Safety and Performance Outcome Measures	18

Table 9. Subject Inclusion / Exclusion Criteria	22

Table 10. State of the Art Inclusion / Exclusion Criteria	23

Table 11. Appraisal Criteria for Suitability and Data Contribution	24

Table 12. Criteria for Appraisal of Literature Level of Evidence	26

Table 13:  Axium Models for US and Australia Markets	30

Table 14:  Preliminary Search Terms	33

Table 15: Publication Inclusion and Exclusion Criteria	34

Table 16: Outcomes Collected	35

Table 17: Literature Review Searches	37

# Table of Figures

Figure 1. Illustration of TNS and INS Components (For Component Identification Only)	11

# Acronyms and Abbreviations

| Term | Definition |
| --- | --- |
| AIMDD | Active Implantable Medical Device Directive |
| AIMD | Active Implantable Medical Device |
| CEP | Clinical Evaluation Plan |
| CER | Clinical Evaluation Report |
| CRPS | Complex Regional Pain Syndrome |
| DRG | Dorsal Root Ganglion |
| EU | European Union |
| FSCA | Field safety corrective actions |
| IFU | Instructions for Use |
| INS | Implantable Neurostimulator |
| MAUDE | Manufacturer and User Facility Device Experience |
| MDD | Medical Device Directive |
| NS | Neurostimulator System |
| PAS | Post Approval Study |
| PMCF | Post Market Clinical Follow-up |
| PNS | Peripheral Nerve Stimulation |
| QoL | Quality of Life |
| SCS | Spinal Cord Stimulation |
| SMNS | Spinal Modulation Neurostimulator System |
| TNS | Trial Neurostimulator |

# Scope of Clinical Evaluation Plan (Stage 0)

This clinical evaluation plan (CEP) summarizes the plan to continuously identify, collect, appraise and analyze relevant clinical data to provide clinical evidence in the Clinical Evaluation Report (CER) for the Abbott Axium Neurostimulator System to demonstrate  conformity with the essential requirements (ERs) for clinical safety and performance when used according to the manufacturer’s Instruction for Use (IFU).

The CEP discusses the plans to address the following stages of the clinical evaluation process:

Stage 0: Define the scope of the clinical evaluation

Stage 1: Identify pertinent data

Stage 2: Appraise pertinent data

Stage 3: Analyze the clinical data

Stage 4: Summarize the clinical evaluation.

The clinical evaluation process is designed to meet the essential requirements as described in AIMDD 90/385/EEC, MEDDEV 2.7/1, Rev 4,  and EN ISO 14155:2011.

## Objective

## Device(s) under Evaluation

The current clinical evaluation includes assessment of current CE-marked, commercially marketed Axium Neurostimulator System (“Axium”) models in EU as well as rationalized Axium products.   Models registered and/or commercially marketed in other geographies such as US and Australia are included in Appendix A.

•	This plan applies to the following products in Table 1 and Table 2. The models currently certified under AIMDD do not have BUDI determinations. Hence, the BUDI and UDI-DI are not applicable (N/A) at this time.

•	Slim Tip Leads: lead model numbers ending with an “A” suffix

•	Ball Tip Leads: lead model numbers with no suffix (no longer manufactured)

Table 1: Axium Products (Current AIMDD Certified)

| Device Name | Model | Device Class , Risk class | Basic UDI | UDI-DI |
| --- | --- | --- | --- | --- |
| IPG Programmers and Accessories |  |  |  |  |
| Clinical Programmer | MN20700 | AIMD | N/A | N/A |
| Patient Programmer | MN20600-02 | AIMD | N/A | N/A |
| Auxiliary Magnet Kit | MN23300 | AIMD | N/A | N/A |
| IPG Chargers and Accessories |  |  |  |  |
| Clinical Programmer Charger Kit | MN23400 | AIMD | N/A | N/A |
| Patient Programmer Charger | MN23400-U | AIMD | N/A | N/A |
| Clinical Programmer Carrying Case | MN23500 | AIMD | N/A | N/A |
| Patient Programmer Carrying Case | MN23500-S | AIMD | N/A | N/A |
| Leads & Lead Accessories |  |  |  |  |
| Implantable SlimTip Lead, 50cm and 90cm* | MN20450-50A, MN20450-90A | AIMD | N/A | N/A |
| Lead Extension Kit* | MN20550-50 | AIMD | N/A | N/A |
| 22cm Small Curve Delivery Sheath Kit* | MN22150 | AIMD | N/A | N/A |
| 22cm Big Curve Delivery Sheath Kit* | MN23650 | AIMD | N/A | N/A |
| Curved Needle* | MN24000 | AIMD | N/A | N/A |
| DRG Lead Accessories Kit* | MN22050 | AIMD | N/A | N/A |
| DRG Connector Cable Kit* | MN21350 | AIMD | N/A | N/A |

*These models were exclusively used with the Axium Neuromodulation System until 2016, when Proclaim DRG was initially released for market distribution. Between 2016 and 2018, their usage have been with Proclaim DRG and Axium Systems. Axium devices manufactured by Spinal Modulation ceased being manufactured in December 2017. These models are covered in the Proclaim DRG CER 90222654 where their usage after 2016 is surveilled.

Table 2: Axium Products (Rationalized from EU)

| Device Name | Models** |
| --- | --- |
| IPG and EPG |  |
| Implantable Neurostimulator | MN20200 |
| Trial Neurostimulator | MN20100 |
| Leads & Lead Accessories |  |
| Implantable Lead (ball tip), 50cm and 90cm | MN20450-50 MN20450-90 |
| Trial Lead (ball tip), 50cm and 90cm | MN20350-50 MN20350-90 |
| Trial SlimTip Lead Kit, 50cm and 90cm | MN20350-50A MN20350-90A |
| Tunneling Tool Kit 30cm/51cm | MN21900 MN22200 |
| 22cm Axium Small Curve Delivery Sheath Kit | MN23850 |
| 22cm Axium Big Curve Delivery Sheath Kit | MN23950 |
| Programmer& Programmer Accessories |  |
| Patient Programmer | MN20600 |

** Axium products which are not marketed or sold in EU.

## Type of Assessment

The clinical evaluation is a routine (regularly scheduled) annual update to assess the clinical safety and performance of the device to demonstrate conformity with the relevant ERs 1, 2 and 5 per AIMDD. . The CER will include an evaluation of the safety, performance, risk/benefit profile, and acceptability of undesirable side-effects as identified through a review of the following:

- Description of device(s) and any design features of the device(s), or any indications or target populations that require specific attention.  This also covers any design features that pose special performance or safety concerns, the intended purpose and application of the device(s) (e.g. target treatment group and disease, proposed warnings, contraindications, precautions, and method of application) and the specific claims made by the manufacturer about the clinical performance and clinical safety of the subject device(s);

- Data from risk management documents. This clinical evaluation will assess the significance of any residual clinical risks that remain after design risk mitigation strategies have been employed;

- The current knowledge/state of the art in the corresponding medical field, such as applicable standards and guidance documents, information relating to the medical condition managed with the device and other treatment alternatives available to the target population;

- Relevant pre-clinical data from verification and validation studies, and clinical data on safety and performance collected through Abbott sponsored pre- and/or post-market clinical studies, peer-reviewed published literature, and post-market surveillance data as applicable. Both favorable and unfavorable data shall be taken into account. Pre-clinical data from verification and validation studies does not apply in this case. Also, no new validation studies have been needed or performed since the last clinical evaluation.

## Timeframe of Assessment

<Indicate time period (X year[s] and date-range if available for literature search(s) and PMS data to be employed in the current evaluation. Note that the date ranges and time period for evaluation may or may not be the same for these inputs. Provide reason or justification for time frames or differences in time frames used if applicable>

Sample text:

The timeframe and period of the current assessment is noted below.

Table 2. Timeframe of Assessment

| Category | TimeFrame | Date Range | Justification |
| --- | --- | --- | --- |
| Lit Search | 1 Year Or 5-years Or Other | XX-XX-XX (date/month/year) -. XX-XX-XX (date/month/year). | This is an annual update hence a 1-year timeframe is selected for evaluation. The date-range for this evaluation is selected to ensure there are no gaps in data analyzed relative to the previous (year XXXX) evaluation (RPTXXXXXX) Or A 5-year timeframe is deemed sufficient to identify relevant and sufficient clinical evidence on the subject device. The date-range for this evaluation is selected to ensure there are no gaps in data analyzed relative to the previous (year XXXX) evaluation (RPTXXXXXX) Or Any other justification as needed/as applicable |
| PMS Data | 5 years Or Other | XX-XX-XX (date/month/year) -. XX-XX-XX (date/month/year). | A 5-year timeframe is used to enable appropriate data for trending and complaint analysis. The date-range for this evaluation is selected to ensure there are no gaps in data analyzed relative to the previous (year XXXX) evaluation (RPTXXXXXX) Or Any other justification as needed/as applicable |

## Newly Emerged Clinical Concerns

The CER update is provided for the purposes of annual update and is not the result of newly emerged clinical concerns.

## Summary of Prior Clinical Evaluation Activities

This CEP is for the update to the existing CER on Axium, and will be Clinical Evaluation Report  CR002 Rev Q. CER revision history is shown in Table 3 below.

Table 3. Axium CER Revision History

| Revision | Description |
| --- | --- |
| A | Initial Release. |
| B | Before: A revision history was provided, but no history overview was included. After: A History Overview has been added to the Appendix to provide ‘Before’ and ‘After’ history for each change. Page numbers were updated. |
| C | Before: No indication language for Nevro was included. Appendix F did not contain a column explaining similarities/differences between devices. No post-market surveillance plan or clinical study reports were included. Need for an additional literature review and conclusion by an impartial expert. After: Appendix F has been updated to be more comprehensive. Indications of Nevro were added and a more detailed comparison of the indications for different SCS companies was added. A post-market surveillance plan and clinical study report summarizing the current clinical trial data were included. Added letters of support from Dr. Timothy Deer and Dr. Liong Liem. |
| D | Add signature sheet to front of report, remove discussion of off-label use in patient registry patients, add units of measure for Charge and Charge Density in Appendix F. |
| E | Change post-market follow-up duration from 3 months to 30 days. |
| F | Add post-market study performance, safety, and effectiveness data; update the Instructions for Use; update the risk analysis data; update format to MEDDEV 2.7.1 r 3, Appendix E; Remove Appendices I, J and K; add new appendices I, J and K; Update Appendix C with DRG publications from Spinal Modulation Inc. |
| G | Update literature search, added dorsal column stimulation/DCS to search terms in the Spinal Cord Stimulation section, updated appendices with new lit search data, updated references with new references, added new CV for lit search personnel. Added new US indication language. Updated to current component configurations and materials. Updated complaints table (thru Q3 2013). |
| H | Update literature search, updated appendices with new literature search data tables, updated SCS/PNS/DRG reference section with new references. Deleted “Ongoing EU Studies” appendix and moved into text (Table 15). Added CV of lit search personnel. Deleted Appendix G (of the previous draft): History Overview. Renumbered appendices. Added SlimTip performance data (Appendix H), DRG bibliography through 04/07/2015 (Appendix K) and vigilance reports (Appendix L). |
| J | Update literature search, updated appendices with new literature search data tables, updated SCS/PNS/DRG reference section with new references, added new CV for literature search personnel. Updated complaints table (Appendix J) and vigilance reports (Appendix L). Added sales numbers for the product (Appendix M). |
| K | Updated literature, updated appendices with new literature search data tables, updated SCS/PNS/DRG reference sections with new references, added new CV for literature search personnel. Updated complaints table (Appendix J) and vigilance reports (Appendix L). Added sales numbers for the product (Appendix M). Updated PMCF reference. Updated Table 6. Updated cardiac implantable device contraindication. |
| L | Added information for US models to Appendices J, L, and M. |
| M | Generalized “Intended Therapeutic Indications” and “Contraindications” statements; updated status of post-market studies in Table 6. |
| N | Updated CER to be compliant with MEDDEV 2.7.1, Rev. 4. |
| P | 2019 Annual Clinical evaluation Update; updated literature from November 1, 2017-December 15, 2018 and post market surveillance data including complaints, FSCA, and CAPAs through November 2018. |
| Q | 2020 Annual Clinical evaluation update, update literature from 01 December 2018 – 31 December 2019 and post market surveillance data including complaints, FSCA and CAPAs through December 31, 2019 |

## Relevant <Essential Requirements> OR <General Safety and Performance Requirements>

Essential Requirements (ERs 1, 2, and 5)  will be assessed using the ER Work Instruction (56-0535) and supported from relevant clinical data in this clinical evaluation. The Essential Requirements Checklist(s) for each product in scope of this evaluation satisfy will be reviewed and evaluated for continued conformance to the Essential Requirements per AIMDD (90506719).

## Applicable Regulatory References

Table 4. Applicable Regulatory References

| Standards Organization and Reference Number | Title |
| --- | --- |
| MDD 93/42/EEC | Council Directive 93/42/EEC of 14 June 1993 Concerning Medical Devices |
| AIMDD 90/385/EEC | Council Directive 90/385/EEC; Annex 7, updated per the 2007/47/EEC M5 amendment, concerning Active Implantable Medical Devices |
| SR No. 236 of 2002 | Therapeutic Goods (Medical Devices) Regulations 2002 |
| MDR 2017/745 | REGULATION (EU) 2017/745 OF THE EUROPEAN PARLIAMENT AND OF THE COUNCIL of 5 April 2017 on medical devices, amending Directive 2001/83/EC, Regulation (EC) No 178/2002 and Regulation (EC) No 1223/2009 and repealing Council Directives 90/385/EEC and 93/42/EEC |
| MEDDEV 2.7/1 Rev 4 | European Commission Guidelines on Medical Devices: Clinical Evaluation: A Guide for Manufacturers and Notified Bodies under Directives 93/42/EEC and 90/385/EEC |
| MEDDEV 2.7/1 Appendix 1 | Appendix 1: Clinical Evaluation of Coronary Stents |
| EMEA/CHMP/EWP/110540/2007 | Guideline on the Clinical and Non-Clinical Evaluation During the Consultation Procedure on Medicinal Substances Contained in Drug-Eluting Coronary Stents |
| IMDRF MDCE WG/N56 Final:2019 | International Medical device Regulator’s Forum Medical device Clinical Evaluation Working Group’s Guidance on Clinical Evaluation |
| ISO 13485:2016 | Medical devices -- Quality Management Systems -- Requirements for Regulatory Purposes |
| EN ISO 14155:2020 | Clinical Investigation of Medical Devices for Human Subjects Good Clinical Practice |
| EN ISO 14971:2019 | Medical devices - Application of Risk Management to Medical Devices |
| Commission Regulation 722/2012 | Commission Regulation 722/2012 Concerning Particular Requirements as Regards the Requirements Laid Down in Council Directives 90/385/EEC and 93/42/EEC with Respect to Active Implantable Medical Devices and Medical Devices Manufactured Utilizing Tissues of Animal Origin |

# Clinical Background and State of the Art

The CER will include a description of the clinical background, treatment options and current knowledge /State-of-the-Art on the treatment of <include Indication or Intended Use verbiage from subject device IFU>.

Comprehensive, objective/non-biased literature searches will be conducted to identify relevant State-of-the-Art literature. Recent reviews and meta-analysis/systematic reviews identified through these searches may also be leveraged as data-leads to identify relevant literature through back-referencing.

Relevant guidelines on the disease state/medical conditions/patient population/ treated by the subject device and clinical literature on the Standard-of-Care/alternative treatment options/competitors/comparators on the subject device itself will be identified through the literature search.

The CER will include a discussion on the Standard of Care/ Relevant benchmark guidelines and relevant clinical evidence on the subject device and competitors/comparators/ alternative treatment options.

Data on safety and performance of key/relevant comparators/competitors will be used to characterize the placement of the subject device in the state of the art. Key comparators/competitors considered for this purpose will include <include names of devices or treatment option categories if already determined/as applicable>

## Literature Search Summary

### Background

A systematic search of the literature will be conducted to identify clinical data relevant to the state of the art regarding the management of chronic intractable pain, and relevant to the safety and performance of the Axium. The review of published literature will describe the clinical background and current knowledge in the relevant medical fields and disciplines relating to the Axium. Data on the safety and performance, risk/benefit profile, and acceptability of undesirable side-effects of the management of chronic intractable pain will be used to characterize the state of the art. Medical alternatives to the management of chronic intractable pain will be evaluated, to determine if alternatives with lower risks or more extensive benefits are available. This data will be used to inform on what is currently considered to provide a high level of protection of health and safety, and what are considered acceptable side effects. Special attention will be given to the recognition and identification of potential new risks, including hazards due to substances and technologies.

### Objectives

The objective of the literature search is to identify clinical data from the scientific literature informing on the safety and performance of Axium and to identify literature informing on the state of the art regarding the management of chronic intractable pain.

To identify information, including both positive and negative findings, relevant for the clinical evaluation of the devices under assessment in the CER, key questions include:

- What are the proper safety and performance outcome measures associated with the use of the Axium?

- What are the clinical results associated with the use of the Axium

- What are the complications or adverse events associated with the use of the Axium?

Literature searches targeting device name keywords will be conducted regarding the following subject to support the evaluation of safety and performance. See Appendix B: Literature Search Protocol for the literature search protocol including literature search strings.

The following assumptions are made:

- Keyword search of peer-reviewed journals through electronic databases will provide required information to address literature review questions.

- Keyword search of peer-reviewed journals through electronic databases will provide valid scientific evidence to address the literature review questions.

- Keywords selected in conjunction with other review methods will adequately capture relevant clinical literature for the review.

### Methods

Using the device name and competitor device name keywords, the Boolean search strings to search multiple databases for relevant literature will be developed in collaboration with an qualified information scientist.

Parameters for the literature searches will be defined as the following:

- Searches targeting the subject devices will be conducted from 01 December 2018 –31 December 2019

The information scientist will execute the searches per the protocol and compile a report of the literature citations obtained from all databases for screening. The literature citations resulting from the searches will be reviewed using a two-pass process. During the first pass, the titles and abstracts will be reviewed to identify potentially relevant articles. During the second pass, full-text articles of citations identified in the first pass will be obtained and reviewed. Articles considered for inclusion in the analysis of safety and performance of Axium will be screened against the inclusion and exclusion criteria in Table 9.

Table 9. Subject Inclusion / Exclusion Criteria

| Exclusion Criteria | Inclusion Criteria |
| --- | --- |
| (E1) The article does not report clinical outcomes relevant to the safety and performance in humans of the Axium Neurostimulator systems (E2) Commentary, review articles, methods papers, editorials, and letters without original data (E4) Not published in English (E5) Case reports with no safety data (E6) Preclinical studies (E7) Multiple devices used from different manufacturers and safety and performance results reported are not device-specific (E8) Duplicated references (E9) Conference abstracts | (I1) Articles reporting clinical outcomes on safety and performance of the Axium Neurostimulator systems when used as intended, with information about device safety and/or performance (I2) Case reports presenting data on adverse events related to the use of the Axium Neurostimulator systems |

Articles considered for inclusion in the analysis of the state of the art will be screened against the inclusion and exclusion criteria in Table 10.

Table 10. State of the Art Inclusion / Exclusion Criteria

| Exclusion Criteria | Inclusion Criteria |
| --- | --- |
| (E1) The article does not inform on the state of the art regarding Neuromodulation, Neurostimulation, Dorsal Root Ganglion Stimulation, DRG Stimulation , the relevant patient population, or treatment alternatives OR report clinical outcomes of the benchmark competitor devices with information about device safety and/or performance relevant to the state of the art (E2) Article presents information covered in another included source (duplicate information) (E3) Commentary, methods papers, editorials, and letters without original data (E4) Not published in English (E5) Case reports without safety data (E6) Preclinical studies (E7) Duplicated references (E8) Conference abstracts | (I1) Articles informing on the state of the art regarding <<medical therapy/device type>>, the relevant patient population, or treatment alternatives (I2) Articles reporting clinical outcomes of the benchmark competitor devices with information about device safety and/or performance relevant to the state of the art |

Articles excluded from the clinical evaluation will be documented with the reason for exclusion in the CER, CR002 Rev Q.

### Results

<Describe the outputs of the literature search and selected references from screening.>

<If appropriate, this section may be moved to an appendix.>

<Provide a brief summary of the literature search strategy for retrieval of information regarding current knowledge and state of the art for the relevant medical fields, medical conditions, and therapy areas. Include mention of key standards and guidance documents applicable to the proposed Indications for Use of the device under evaluation or the device itself.>

<Recommendations from clinical practice guidelines and consensus statements published by global heart rhythm and cardiology societies serve as the primary source of information in this discussion. Guidance and consensus documents reflect the judgment of acknowledged subject matter experts, are based on their experience as well as a detailed examination of the available literature, representing the current thinking in the field.>

<<<Insert summary of the process and outcomes of search process>>>

Table 5. Search Terms

Dialog Databases: Allied & Complementary Medicine™, Analytical Abstracts, Embase®, EMCare®,  International Pharmaceutical Abstracts,  MEDLINE®,  ToxFile®

| Set# | Searched for | Results |
| --- | --- | --- |
| S4 | <<insert search string>> | <<number>> |
| S3 | <<insert search string>> | <<number>> |
| S2 | <<insert search string>> | <<number>> |
| S1 | <<insert search string>> | <<number>> |

<Sample figure for search process and results—Summarizing the number of publications that resulted and were reviewed against the selection criteria; number of articles that met the selection criteria, any additional references that were identified in the subsequent review of literature>

Figure 1. State of the Art Literature Review Selection Process and Results

| Identification | Records identified through database searchings N = 385 |  | Additional records identified through other sources N = 75 |  |
| --- | --- | --- | --- | --- |
|  | Records after duplicates removed N = 459 |  |  |  |
| Screening | Records screened N = 459 |  | Records excluded N = 174 -------------------------------------------------------- (E1) No abstract available (except clinical pra... (N = 38) (E3) Duplicated references (N = 23) (E5) Commentary, methods papers, editorials, an... (N = 1) (E6) Article presents information covered in an... (N = 4) (E7) Article does not meet any inclusion criteria (N = 108) |  |
| Eligibility | Full-text articles assessed for eligibility N = 285 |  | Full-text articles excluded, with reasons N = 206 ------------------------------------------------------------------------ (E3) Duplicated references (N = 58) (E4) Conference abstracts (N = 7) (E6) Article presents information covered in an... (N = 15) (E7) Article does not meet any inclusion criteria (N = 126) |  |
| Included | Studies included N = 79 |  |  |  |

## Literature Review

### Medical Conditions

<Describe EACH medical field concerned/relevant medical condition or disease to be treated, managed, or diagnosed with the subject device(s), which should align the intended purpose/medical indications in the IFU. Create subsections as necessary.>

<Describe natural course and consequences of the medical conditions concerned>.

- Information on the clinical condition(s) to be treated, managed, or diagnosed

- Prevalence of the condition(s)

- Natural course of the condition(s)

<Describe the types of users of the device under evaluation, their needs, and their unmet needs in relation to treating the medical conditions concerned>.

### Technological Developments

<Introduce subject device category and explain how the subject device supports the target therapy:

- Explain how the technology used for the target therapy has evolved over time (the degree of maturity as a technology)

- Detail the various types of medical devices used with/supporting the target therapy including the subject device

- Mechanism of action

- Design characteristics affecting clinical performance

- Describe and compare key features and specifications of subject device and competitor devices if applicable

- Compare safety and performance/clinical outcomes of the medical devices (subject device and competitor devices) under the target therapy by using data from the literature:

- Use head-to-head comparative studies (preferred)

- For each study cited, define study type (RCT, multi-center, controlled, double-blind, etc.)

- Present side-by-side tables of clinical outcomes with subject and competitor devices as applicable (cited from literature); if not head-to-head studies, include relevant information on trial design, study population, outcome assessment such that heterogeneity can be recognized

- Present side-by-side tables of AEs (cited from literature) with subject and competitor devices as applicable (using same method as above)

- If many studies available and not all presented, identify selection criteria (studies >x patients; RCT only; studies published since YEAR) and justification (most informative; studies with current design iterations)

- Note: This information may not be available for devices considered ancillary to procedures and thus rarely identified (guide wires, catheters, accessories). In this case, refer to target intervention performance and safety data listed in section 3.3. For some devices, device-specific safety information can be extracted from the literature as certain procedural AE can be linked to certain devices (example: guide wire/endovascular intervention, subintimal tracking and dissection)

- Conclude which is the current role of the subject device. The position within the treatment portfolio has to be specified properly in the clinical evaluation report and other relevant documentation.

- The subject device(s) when used for target therapy can/cannot compete with an/multiple agreed first-line treatments/best in class

- The subject device(s) may add to the range of acceptable treatments, as even a first-line treatment will likely have contraindications or non-responders

- The subject device(s) may also be supported by clinical evidence for an acceptable benefit/risk-profile for specific, defined subgroups

- The subject device(s) may be supported for use under specific conditions (e.g. emergency situations)

### Therapeutic, Diagnostic, or Treatment Options Summary

<Describe descriptions of available therapeutic, management, and diagnostic options.>

<Include historical context and developments, summary of advantages and disadvantages of various options, risk-to-benefit profile for various options, and limitations of various options.>

- available therapeutic, management, and diagnostic options. (e.g., other devices, medical alternatives including conservative, surgical and medicinal)

- historical context and developments,

- summary of advantages and disadvantages of various options,

- risk-to-benefit profile for various options, and

- limitations in relation to the different clinical forms, stages, and severities of the medical condition and in relation to different target populations.

- benefit and risks (nature, extent, probability, duration, frequency),

- acceptability of undesirable side-effects and other risks (including the nature, severity, probability and duration of acceptable harm.

- hazards due to substances or technologies that could be relevant to the device under evaluation

- Mechanisms of harm, clinical aspects of minimization and management of side effects and other risks.

- Types of users

- Diverging opinions of professionals as to the use of different medical options

- Unmet medical needs.

<Describe any hazards due to substances or technologies that could be relevant to the device under evaluation.>

### Similar Devices

<Include a list of relevant similar and competitive devices, including Abbott devices. >

<Summarize the literature regarding benchmark or pivotal trials for the similar devices with a focus on safety and performance measures that will be used for comparison with the subject device. Use the appraisal process to limit article selection to those with the highest level of evidence, if appropriate.>

<The following sample table can be used to present a summary of the safety and performance outcomes for the competitive devices to be used to compare with the subject devices.>

Table 6. Competitive Device Literature Summary

| Study | Study Design | Device Application | Performance Outcomes | Safety Outcomes |
| --- | --- | --- | --- | --- |

### State of the Art Landscape Analysis

<List/describe benchmark/available similar devices (available in Europe) and key properties (device, manufacturer, key design features, specifications, mechanism of action, use/indication as applicable).>

<Comparative analysis of subject device vs similar device performance and safety:

- Head-to-head trials (safety /performance)

- Side-by side tables: Clinical outcomes (cited from literature)

- Side-by side tables: Clinical safety (cited from literature)

<Use Table 7 to summarize the comparative analysis.>

Table 7. Published Studies of <Subject/Equivalent Device Name> and Similar Devices

| Article | Setting | Arms, No Patients | Procedural Outcomes | Safety Outcomes | Performance/Other |
| --- | --- | --- | --- | --- | --- |
| <EndNote in-text citation> | <Intervention> <Key patient characteristics> <Location if relevant> | <Device A>, n=xx vs <Device B>, n=xx | <Device A vs Device B> <Endpoint 1>, n (%) vs n (%); P=.xx <Endpoint 2>, n (%) vs n (%); P=.xx | <Device A vs Device B> <Endpoint 1>, n (%) vs n (%); P=.xx <Endpoint 2>, n (%) vs n (%); P=.xx | <Relevant subgroup analyses> <Relevant information on inclusion criteria/procedure characteristics that may influence outcomes> |

### Conclusion

< Provide a conclusion of the literature search summary, literature review of medication conditions, technological developments/similar devices, SOTA landscape and placement of the device within the SOTA>

Sample text (for conclusion if applicable):

Based on this review of published guidelines and clinical data, it is demonstrated that the <<<Insert Device Name(s)>>> under assessment in this CER <<<continue to meet or continue>>> state of the art criteria in terms of therapeutic appropriateness for their intended use, feature availability, clinical benefit/risk profiles, and clinical safety and performance.

# Safety and Performance Measures/Objectives

# Device Description

## Detailed Device Description

The Axium Neurostimulator System is composed of two subsystems as pictured in Figure 1. The Axium Neurostimulator System is a fully-implantable medical device that stimulates neural tissue to produce analgesia for the treatment of chronic pain.

Figure 1. Illustration of TNS and INS Components (For Component Identification Only)

The following is a description of each of the components of the Axium Neurostimulator System:

- Trial Neurostimulator (TNS)

The external TNS device connects to the Trial Lead(s) or Lead Extensions and is worn by the patient for up to 30 days during the trial period. The TNS device has a belt clip for the patient’s convenience.

- Implantable Neurostimulator (INS)

The Axium Implantable Neurostimulator (INS) is a non-rechargeable, 4-channel electronic device. It uses microelectronic circuitry, powered by a hermetically sealed battery, to generate a pulsed waveform to stimulate neural tissue. The electronic circuitry and battery are housed in a hermetically sealed titanium case.

Each neurostimulator has a unique internal identifier that allows the physician to identify the type of device through an X-ray. The radiopaque identifier inside the case allows identification of both the device manufacturer and model number using standard x-ray equipment. For the Axium Neurostimulator, the code is SM001, which identifies Spinal Modulation as the manufacturer.

The INS is packaged in a sealed inner tray within a sealed outer tray.

- Implant Leads/Trial Leads/Lead Extension

The Lead Kits contain the Leads and the individual delivery devices that are required for their placement.

Implant / Trial Leads: The Leads are designed for percutaneous introduction into the body using a special needle and a set of custom delivery tools provided in their respective kits. Each Lead is fitted with four cylindrical electrodes spaced at equal intervals, which are intended to provide stimulation at the target dorsal root ganglion (DRG). The temporarily placed Trial Leads are intended for use for up to 30 days.

Lead Extension: The Lead Extension consists of a silicone port header that accepts the Axium Trial Lead and Implant Lead. It is intended to extend the length of the lead and provide a connection between the lead and the Connector Cable, or the lead and the Implantable Neurostimulator System. The Lead Extension is intended for chronic implantation as a component of the Axium Neurostimulator System.

- Connector Cable Kit

The Connector Cable connects the Leads or Lead Extension to the external TNS. The Connector Cable is packaged separately from the Lead and Lead Extension Kit. The Connector Cable includes a connector and 2 extension cables for use as needed.

The Lead Identifiers are small bent wires that can be inserted into the proximal stylet lumen in the leads and are individually colored to identify the leads.

- Lead Accessories

Small / Big Curve Delivery Sheaths are intended to allow passage of the lead percutaneously into the epidural space. The labeled length of the sheath is the distance from the hub to the pre-shaped tip and the length of the curve at the tip is approximately 2 mm for the Small Curve and approximately 8mm for the Big Curve.

Axium Small / Big Curve Delivery Sheaths are also intended to allow passage of leads percutaneously into the epidural space. Axium sheaths are internally reinforced with thin stainless steel braiding.

Complex Curve and Straight Stylets assist in steering and positioning the lead within the epidural space. These are included within the 50cm and 90cm SlimTip trial lead kits.

The 14G Delivery Needle is intended to access the epidural space, providing a conduit for lead, guidewire and delivery sheath placement. It is available as a straight needle or a curved needle. The Axium Curved Needle is available in a separate package and is uniquely identified by the colored hub. The 14G Delivery Needle is only available in the Implant Lead, Trial Lead, and Lead Accessories Kits.

The Guidewire is optional and is intended to verify that the needle is in the epidural space after using a loss of resistance technique.

Soft Tissue Anchors are intended to anchor the Lead in the soft tissue or on the skin surface proximal to the distal contacts of the Lead.  The soft tissues anchors are packaged with the leads, and included in the lead accessories kit.

- Implantation Tools

The tunneling tool is used to provide a conduit for the Trial Lead, Implant Lead, or Lead Extension to the INS or away from the midline of the spine. It is packaged with 2 exchangeable tips: a blunt pencil tip and a sharp trocar tip. A straw is slid over the tunneling tool and when the steel handle is removed, the straw provides the conduit for tunneling.

The INS Sizer is approximately the same size as the INS and allows the physician to properly size the INS pocket. The INS sizer is packaged with the Axium tunneling tool kits (MN21900, MN22200).

The port plugs are used to fill unused ports in the INS. The port plugs are packaged with the INS, but spare port plugs are also packaged with the Tunneling Tool Kits for the convenience of the physician.

- Additional Accessories

Sterile Magnet Sleeve: The magnet is placed in the sterile sleeve to allow it to be used during the implantation of the INS. The sterile magnet sleeve is packaged with the Axium tunneling tool kits (MN21900, MN22200).

Medical Alert Card identifies the patient as a user of the Neurostimulation System. The Medical Alert card is packaged with the INS and Patient Programmer.

Programmer Charger is to be used with the Clinical or Patient Programmers to charge the battery or allow use of the Programmers while plugged into standard electrical outlets.

Programmer Carrying Case protects the Programmers from liquid.

Auxiliary Magnet allows the user to turn the NS off or activates RF to allow the user to communicate with the NS.

Hex Key allows the user to release a set screw in the INS header or Lead Extension header that has been unscrewed too far. The Hex Key is not to be used to tighten the set screw against the lead. The hex key is packaged with the Axium tunneling tool kits (MN21900, MN22200).

Clinical Programmer and Patient Programmer

The Clinical Programmer is used to program the stimulation parameters for both the TNS and the INS. The instructions for programming the TNS and INS devices are the same. The Clinical Programmer is used by the physician or clinical staff.

The Patient Programmer allows the patient to adjust the stimulation settings of the TNS and INS devices within limits preset by the clinician. The Patient Programmer also allows the patient to turn stimulation off, if necessary.

Sample Text/table option

Table 10 provides subject device specifications based on model.

Table 10. Subject Device Specifications

| Model Number | Size | Usable Length | Curl Type | Curl Dimension | Dilator (OD × L) | Guidewire OD × L, Tip |
| --- | --- | --- | --- | --- | --- | --- |

## Name and Address of Manufacturer

Sample text:

The legal manufacturer and manufacturing site (s) are listed in Table 11.

Table 11. Name and Address of Manufacturer

| Device | Legal Manufacturer | Manufacturing Site |
| --- | --- | --- |

## Principles of Operation

<Provide description of application and mode of action>

## Development History

<Include timing and nature of design changes/evolution during the course of the Device lifetime leading up to the current evaluation.  Include names, models, sizes of the predecessor device, whether the predecessor device is still on the market, description of the modifications.>

<Discuss if the clinical impacts of the design evolution have been tested /if the changed device is currently on the market, or, the plan for testing these.>

## Regulatory History

<Include a summary of CE mark history.>

Sample text:

The CE Mark approval and certification timelines for the <<<subject devices(s)>>> are shown in Table 12.

Table 12. <<<Insert Device Name(s)>>> Regulatory History

| Device Name | Model Number | GMDN Code | Date of First CE Mark | Current Certificate (optional) |
| --- | --- | --- | --- | --- |

## Product Labeling

<Directly reference text from EU IFU(s) or draft if commercial IFU not available. Not all sub-content noted below may be available within all IFUs If other sources referenced, clearly cite those references and align with relevant stakeholders.. >

### Intended Purpose/Use

This Axium Neurostimulator System is designed to deliver low-intensity electrical impulses to nerve structures. The system is intended to be used with leads and associated extensions that are compatible with the system

### Intended Target Groups

Patients

The Axium Neurostimulator System is intended to treat patients with chronic intractable pain that has been unresponsive to other methods.

Users

Implanting physicians should be experienced in the diagnosis and treatment of chronic pain syndromes and have undergone surgical and device implantation training

### Indications

The Axium Neurostimulator System is indicated for the management of chronic intractable pain

### Contraindications

Patients contraindicated for the Axium Neurostimulator System are those who:

- Have an active implantable medical device including but not limited to cardiac pacemakers and cardiac defibrillators

- Are unable to operate the system

- Are poor surgical risks

Are pregnantAre under the age

### Warnings

<May reference IFU including specific version rather than including all here.>

Sample text:

Reference the subject device IFU (version#) for complete list of warnings.

### Precautions

<May reference IFU including specific version rather than including all here.>

Sample text:

Reference the subject device IFU (version#) for complete list of precautions.

### Potential Adverse Events

<Include a list off all potential adverse events referencing IFU and other source   document(s) as appropriate.  Note that these potential adverse events should also include and be consistent with those in the risk documentation, RML for MDR CERs, literature and PMS summarized in the CER.>

## Intended User

<Summarize the medical professionals intended to use the subject device (s).>

## Intended Clinical Benefits

Previous clinical evaluations show that the Axium DRG neurostimulation therapy benefits patients with chronic intractable pain who have no other options and that subjects are satisfied with the treatment. The safety profile has remained similar to the results from other studies and should be considered as an acceptable risk given these patients have failed multiple therapies before their neurostimulator implant.  Factors such as quality of life and improved sleep are also positively affected by DRG stimulation. The safety profile remains similar to the results from the pre-market study, and should be considered as an acceptable risk given these patients failed multiple therapies before their neurostimulator implant. The current clinical evaluation will assess the Axium system for continued benefit and risk analysis and results will be reported in the CER.

## Claims

<Summarize safety and performance claims associated with manufacturer's subject device IFU and promotional materials for the subject device per MEDDEV 2.7/1 Rev 4.>

<NOTE: Devices such as Class I accessories, which are not patient contacting, there may be no or very limited safety and performance claims. Align with CEC or relevant stakeholders to determine content in this Section for such devices.>

Sample text:

Table 13 summarizes the safety and performance claims that have been established for the subject device(s).

Table 13. Safety and Performance Claims

| Claim | Source | Verification/Substantiation |
| --- | --- | --- |

## Devices Used in Combination (if applicable)

<Identify and include other devices and products, which may be used in combination with the subject device(s) in scope of this clinical evaluation. >

## Animal Tissues or Derivatives (if applicable)

<The information in this section may be obtained from Regulatory.>

Sample text:

There are no components of animal origin in the <subject device. Any chemicals, such as mold release agents considered as tallow derivatives have been processed under conditions at least as rigorous as those specified in ISO 22442-1, Annex C, Section C.5, are considered exempt from Commission Regulation 722/2012 of 08 August 2012 (Article 1 Section 4(a)). These are considered acceptable for Transmissible Spongiform Encephalopathies risk due to their processing conditions and not their sourcing, irrespective of the geographical origin and the nature of the tissues from which tallow derivatives are derived.

<NOTE: For devices, such as Class I accessories which are not patient contacting, this section may not be relevant. Align with relevant stakeholders to determine content in this Section for such devices.>

# Clinical Development Plan (Stage 0)

<This section is an MDR requirement and may be removed if the subject devices remain certified under the MDD/AIMDD.>

<Outline the clinical development plan for the subject device. Reference the CDP created by Clinical Affairs, if applicable. Insert the CDP table, if applicable. If no CDP has been created, involve Clinical Affairs in the development of content for this section.>

Sample text:

The clinical development plan (CDP) provides an overview of the progression from exploratory investigations, such as first-in-man studies, feasibility and pilot studies, to confirmatory investigations, such as pivotal investigations, and a PMCF for the device under evaluation. The CDP (as required by Annex XIV: 1a) for <<Device>> is summarized in Table 14.

Table 14. Clinical Development Plan Summary

| Type of Clinical Evidence | Potential Acceptance Criteria | Milestone | Estimated Timeline | References |
| --- | --- | --- | --- | --- |
| (e.g. IDE trial, PMCF trial, physician survey, etc) | Include clinical safety, clinical performance, and clinical benefit assessments. | (e.g. Complete enrollments, Complete final report, etc) | Included estimated complete dates (e.g. Dec 2018) | Refer to CIP or CIR numbers |

# Equivalency Assessment (Stage 0)

Equivalence to other devices will not be claimed for the Axium Neurostimulator System.

# 6.1	Identify Pertinent Data for Clinical Evaluation (stage 1)

The CER will include an evaluation of the safety, performance, risk/benefit profile and acceptability of undesirable side-effects of the Axium Neurostimulator System as identified through a review of the following data sources and types of data:

- Safety and performance data collected through company sponsored Pre- and/or Post-market clinical studies of the Axium Neurostimulator System;

- Post-market surveillance data from the Axium Neurostimulator System;

- Published literature describing the current knowledge and state of the art therapies for the indications for use of these spinal cord stimulator devices;

- Systematic review of recently published peer-reviewed literature clinically relevant to the Axium Neurostimulator System and equivalent market released devices.

## Equivalent Device(s)

<Identify the equivalent device(s) and legal manufacturer(s). If the equivalent device is a competitor device, provide evidence that a current valid contract is in place between the Abbott and the device manufacturer allowing ongoing access to the technical documentation. If equivalence is claimed to more than one device, each demonstration of equivalence can only be based on a single device. Each equivalence device must meet all three equivalence criteria (clinical, technical, biological).>

## Equivalence Tables

<The equivalence tables below are taken from MDCG 2020-5. Use the tables to fully identify and disclose any differences between the two devices, and to provide the scientific justification for equivalence. Include references to pre-clinical data as part of the scientific justification (where relevant). The items in the first column of the table are examples only and are to be considered as such. They must not be interpreted as an exhaustive list of specifications, properties, parameters and/or aspects for demonstrating equivalence to another device.>

<Where more than one device is assessed for equivalence, the table should be completed separately for each presumed equivalent device.>

Table 15. Technical Equivalence Table

| Equivalence Table For the comparison of a device with a presumed equivalent marketed device for the purpose of demonstrating equivalence |  |  |  |
| --- | --- | --- | --- |
| - Technical Characteristics (add a separate row for each of the assessed characteristics) | Device 1 (under clinical evaluation) Description of characteristics and reference to specifying documents | Device 2 (marketed device) Description of characteristics and reference to specifying documents | Identified differences or conclusion that there are no differences in the characteristic |
| Device is of similar design |  |  | 1.1 |
| Used under similar conditions of use |  |  | 1.2 |
| Similar specifications and properties including physicochemical properties such as intensity of energy, tensile strength, viscosity, surface characteristics, wavelength and software algorithms |  |  | 1.3 |
| Uses similar deployment methods where relevant |  |  | 1.4 |
| Has similar principles of operation and critical performance requirements |  |  | 1.5 |

| Scientific justification why there would be no clinically significant difference in the safety and clinical performance of the device, OR a description of the impact on safety and or clinical performance (Use one row for each of the identified differences in characteristics, and add references to documentation as applicable) | Clinically significant difference Yes / No |
| --- | --- |
| 1.1 |  |
| 1.2 |  |
| 1.3 |  |
| 1.4 |  |
| 1.5 |  |

Table 16. Biological Equivalence Table

| Equivalence Table For the comparison of a device with a presumed equivalent marketed device for the purpose of demonstrating equivalence |  |  |  |
| --- | --- | --- | --- |
| - Biological Characteristics (add a separate row for each of the assessed characteristics) | Device 1 (under clinical evaluation) Description of characteristics and reference to specifying documents | Device 2 (marketed device) Description of characteristics and reference to specifying documents | Identified differences or conclusion that there are no differences in the characteristic |
| Uses the same materials or substances in contact with the same human tissues or body fluids |  |  | This characteristic must be the same for the demonstration of equivalence 2.1 |
| Similar kind and duration of contact with the same human tissues or body fluids |  |  | 2.2 |
| Similar release characteristics of substances including degradation products and leachables |  |  | 2.3 |

| Scientific justification why there would be no clinically significant difference in the safety and clinical performance of the device, OR a description of the impact on safety and or clinical performance (Use one row for each of the identified differences in characteristics, and add references to documentation as applicable) | Clinically significant difference Yes / No |
| --- | --- |
| 2.1 |  |
| 2.2 |  |
| 2.3 |  |

Table 17. Clinical Equivalence Table

| Equivalence Table For the comparison of a device with a presumed equivalent marketed device for the purpose of demonstrating equivalence |  |  |  |
| --- | --- | --- | --- |
| - Clinical Characteristics (add a separate row for each of the assessed characteristics) | Device 1 (under clinical evaluation) Description of characteristics and reference to specifying documents | Device 2 (marketed device) Description of characteristics and reference to specifying documents | Identified differences or conclusion that there are no differences in the characteristic |
| Same clinical condition or purpose, including similar severity and stage of disease |  |  | 3.1 |
| Same site in the body |  |  | (The characteristic must be the same for the demonstration of equivalence) 3.2 |
| Similar population, including as regards age, anatomy and physiology |  |  | 3.3 |
| Same kind of user |  |  | (The characteristic must be the same for the demonstration of equivalence) 3.4 |
| Similar relevant critical performance in view of the expected clinical effort for a specific intended purpose |  |  | 3.5 |

| Scientific justification why there would be no clinically significant difference in the safety and clinical performance of the device, OR a description of the impact on safety and or clinical performance (Use one row for each of the identified differences in characteristics, and add references to documentation as applicable) | Clinically significant difference Yes / No |
| --- | --- |
| 3.1 |  |
| 3.2 |  |
| 3.3 |  |
| 3.4 |  |
| 3.5 |  |
| Summary In the circumstance that more than one non-significant difference is identified, provide a justification whether the sum of differences may affect the safety and clinical performance of this device. |  |

# Identify Pertinent Data for Clinical Evaluation (stage 1)

<As an introduction to the section, state the high level sources for the clinical evaluation.>

Sample text:

Data for the clinical evaluation of <<device>> consists of critically-evaluated scientific literature for <<subject or equivalent device>> and current knowledge/state of the art; and post-market clinical experience data including patient-effect complaints and device complaints for <<subject or equivalent device>> was included, as appropriate.

## Data Sources

<Describe/list the specific data sources for the clinical evaluation and provide the location within the CER that details the appraisal of each source.>

Sample text:

Comprehensive methods were utilized to identify clinical and non-clinical data relevant for the clinical evaluation of <<subject device>>. These include non-clinical biocompatibility evaluation and bench-top testing data, data from critically evaluated scientific literature, clinical investigations, current knowledge / SOA; and post-market clinical experience data from patient-effect and device complaints, external vigilance data, post-market surveillance plans, and post-market clinical follow-up. Table 18 lists each data source and the sections detailing their description.

Table 18. CER Data Sources

| Data Source | Section (Detailed Description) |
| --- | --- |
| Non-clinical data (Bench-Top Testing Data, Biocompatibility Testing, Pre-clinical Studies) | Section XX |
| Clinical Investigation Data | Section XX |
| Literature | Section XX |
| PMCF Data | Section XX |
| Customer Complaint Data | Section XX |
| Physician Preference Testing | Section XX |

## Non-Clinical Data

## Clinical Investigations (or Justification for No Clinical Investigation)

## Pre-Clinical Studies

There are no new pre-clinical studies during this clinical evaluation. All pre-clinical studies for the products in scope of this evaluation are covered in the existing CER CR002. All products described in the CER have successfully passed formal verification and validation.

## Pre-Market Clinical Investigations

There are no new pre-market clinical investigations during this clinical evaluation. The completed pre-market clinical investigation studies are covered in CER CR002.

## Post-Market Clinical Follow-Up

Completed Post market studies involving Axium products are summarized in the current CER, CR002, Rev P and will also be included in the Rev Q CER. Updates and any available data from the following active PMCF activities will be analyzed in the CER:

- TARGET (Targeted SCS (DRG) Stimulation) post-approval study (PAS; CRD 810; Protocol SJM-CIP-10113). The TARGET study is a post market study is designed to demonstrate continued safety of spinal cord stimulation of the dorsal root ganglion (DRG) for the treatment of chronic, intractable lower extremity neuropathic pain associated with a diagnosis of complex regional pain syndrome (CRPS) and peripheral causalgia. NOTE: This study is a follow on from the ACCURATE study and includes data for both Axium and Proclaim devices.   Current AIMDD certified Axium devices Auxiliary Magnet Kit (MN23300) and Programmer Charger Kit (MN23400) are in the TARGET study.  The study includes US models of Axium devices.

REALITY (CRD 959; Protocol ABT-CIP-10279 which includes the current AIMDD certified and commercially available Axium Clinician and Patient Programmers and other accessories

## Data Retrieved from Literature

### Literature Search Protocol

#### Background

A systematic search of the literature will be conducted to identify clinical data relevant to the state of the art regarding the management of chronic intractable pain, and relevant to the safety and performance of the Axium. The review of published literature will describe the clinical background and current knowledge in the relevant medical fields and disciplines relating to the Axium. Data on the safety and performance, risk/benefit profile, and acceptability of undesirable side-effects of the management of chronic intractable pain will be used to characterize the state of the art. Medical alternatives to the management of chronic intractable pain will be evaluated, to determine if alternatives with lower risks or more extensive benefits are available. This data will be used to inform on what is currently considered to provide a high level of protection of health and safety, and what are considered acceptable side effects. Special attention will be given to the recognition and identification of potential new risks, including hazards due to substances and technologies.

#### Objective

The objective of the literature search is to identify clinical data from the scientific literature informing on the safety and performance of Axium and to identify literature informing on the state of the art regarding the management of chronic intractable pain.

To identify information, including both positive and negative findings, relevant for the clinical evaluation of the devices under assessment in the CER, key questions include:

- What are the proper safety and performance outcome measures associated with the use of the Axium?

- What are the clinical results associated with the use of the Axium

- What are the complications or adverse events associated with the use of the Axium?

Literature searches targeting device name keywords will be conducted regarding the following subject to support the evaluation of safety and performance. See Appendix B: Literature Search Protocol for the literature search protocol including literature search strings.

The following assumptions are made:

- Keyword search of peer-reviewed journals through electronic databases will provide required information to address literature review questions.

- Keyword search of peer-reviewed journals through electronic databases will provide valid scientific evidence to address the literature review questions.

- Keywords selected in conjunction with other review methods will adequately capture relevant clinical literature for the review.

#### Methods

Using the device name and competitor device name keywords, the Boolean search strings to search multiple databases for relevant literature will be developed in collaboration with an qualified information scientist.

Parameters for the literature searches will be defined as the following:

- Searches targeting the subject devices will be conducted from 01 December 2018 –31 December 2019

The information scientist will execute the searches per the protocol and compile a report of the literature citations obtained from all databases for screening. The literature citations resulting from the searches will be reviewed using a two-pass process. During the first pass, the titles and abstracts will be reviewed to identify potentially relevant articles. During the second pass, full-text articles of citations identified in the first pass will be obtained and reviewed. Articles considered for inclusion in the analysis of safety and performance of Axium will be screened against the inclusion and exclusion criteria in Table 9.

Table 9. Subject Inclusion / Exclusion Criteria

| Exclusion Criteria | Inclusion Criteria |
| --- | --- |
| (E1) The article does not report clinical outcomes relevant to the safety and performance in humans of the Axium Neurostimulator systems (E2) Commentary, review articles, methods papers, editorials, and letters without original data (E4) Not published in English (E5) Case reports with no safety data (E6) Preclinical studies (E7) Multiple devices used from different manufacturers and safety and performance results reported are not device-specific (E8) Duplicated references (E9) Conference abstracts | (I1) Articles reporting clinical outcomes on safety and performance of the Axium Neurostimulator systems when used as intended, with information about device safety and/or performance (I2) Case reports presenting data on adverse events related to the use of the Axium Neurostimulator systems |

Articles considered for inclusion in the analysis of the state of the art will be screened against the inclusion and exclusion criteria in Table 10.

Table 10. State of the Art Inclusion / Exclusion Criteria

| Exclusion Criteria | Inclusion Criteria |
| --- | --- |
| (E1) The article does not inform on the state of the art regarding Neuromodulation, Neurostimulation, Dorsal Root Ganglion Stimulation, DRG Stimulation , the relevant patient population, or treatment alternatives OR report clinical outcomes of the benchmark competitor devices with information about device safety and/or performance relevant to the state of the art (E2) Article presents information covered in another included source (duplicate information) (E3) Commentary, methods papers, editorials, and letters without original data (E4) Not published in English (E5) Case reports without safety data (E6) Preclinical studies (E7) Duplicated references (E8) Conference abstracts | (I1) Articles informing on the state of the art regarding <<medical therapy/device type>>, the relevant patient population, or treatment alternatives (I2) Articles reporting clinical outcomes of the benchmark competitor devices with information about device safety and/or performance relevant to the state of the art |

Articles excluded from the clinical evaluation will be documented with the reason for exclusion in the CER, CR002 Rev Q.

## Post-Market Surveillance Data

Post-market surveillance (PMS) data generated from EU and world-wide complaints and reportable vigilance incidents associated with the Axium received through December 31, 2019 will be analyzed in the CER.

PMS data generated from corrective and preventative actions (CAPAs) and field safety corrective actions (FSCAs) associated with the Axium through December 31, 2019 will be analyzed in the CER

## Post-Market Clinical Follow-Up Data

# Appraisal of Pertinent Data for Clinical Evaluation (STAGE 2)

### Appraisal Criteria

Each of the identified clinical data sets will be appraised for their suitability and data contribution toward the clinical evaluation, according to the Global Harmonization Task Force SG5/N2R8 guideline and MEDDEV 2.7/1 Rev 4. The appraisal criteria are presented in Table 11.

Table 11. Appraisal Criteria for Suitability and Data Contribution

| Appraisal Criteria for Suitability |  |  |
| --- | --- | --- |
| Suitability Criteria | Description | Grading System |
| Appropriate Device | Were the data generated from the device in question? | D1 = Actual device D2 = Comparable device D3 = Other device |
| Appropriate Device Application | Was the device used for the same intended use (e.g., methods of deployment, application, etc.)? | A1 = Same Use A2 = Minor deviation A3 = Major deviation |
| Appropriate Patient Group | Were the data generated from a patient group that is representative of the intended treatment population (e.g., age, sex, etc.) and clinical condition (i.e., disease, including state and severity)? | P1 = Applicable P2 = Limited P3 = Different population |
| Acceptability of the Report / Data Collection | Do the reports or collations of data contain sufficient information to be able to undertake a rational and objective assessment? | R1 = High quality R2 = Minor deficiencies R3 = Insufficient information |
| Appraisal Criteria for Data Contribution |  |  |
| Data Contribution Criteria | Description | Grading System |
| Data Source Type | Was the design of the study appropriate? | T1 = Yes T2 = No |
| Outcome Measures | Do the outcome measures reported reflect the intended performance of the device? | O1 = Yes O2 = No |
| Appropriateness of Follow-up Period | Is the duration of follow-up long enough to assess whether duration of treatment effects and identify complications? | F1 = Yes F2 = No |
| Statistical Significance | Has a statistical analysis of the data been provided and is it appropriate? | S1 = Yes S2 = No |
| Clinical Significance of the Treatment Effect | Was the magnitude of the treatment effect observed clinically significant? | C1 = Yes C2 = No |
| Appraisal criteria defined per guidance in the GHTF SG5/N2R8:2007 Guidance, Appendix D, A Possible Method of Appraisal |  |  |

## Appraisal Criteria for Quality of Data

To assess quality of clinical data, each clinical data set will be appraised in terms of scientific validity, and weighted according to the pre-defined criteria described in Table 12 below. This methodology of appraisal is based on the evaluation of outcomes from use of the subject devices or the similar devices, the study design and sample size.  It also emphasizes data obtained in large studies (>100 patients), utilizing a randomized control design over case series, and prospective studies over retrospective data collection. The data cut-offs with respect to study size are arbitrary, recognizing that event rates derived in series that include less than 30 patients are not as reliable as those obtained in larger series with greater than 100 patients.

Table 12. Criteria for Appraisal of Literature Level of Evidence

| Level of Evidence | Criteria |
| --- | --- |
| A1 | Prospective randomized controlled trials (RCT) comparing outcomes of ≥100 patients <<implanted with, treated with, using>> the subject device (Abbott / SJM Axium Neurostimulator System to control patients |
| A2 | Prospective randomized controlled trials (RCT) comparing outcomes of 30 - 99 patients <<implanted with, treated with, using>> the subject device (Abbott / SJM Axium Neurostimulator System to control patients or Prospective cohort trials comparing outcomes of ≥100 patients <<implanted with, treated with, using>> the subject device (Abbott / SJM Axium Neurostimulator System to a control cohort or Meta-analysis of RCT, cohort or case-series studies or compilation report from user experience databases (e.g., MAUDE, vigilance reporting) evaluating the safety and performance of the subject device (Abbott / SJM Axium Neurostimulator System) |
| B1 | Prospective cohort trials comparing outcomes of 50-99 patients <<implanted with, treated with, using>> the subject device Abbott / SJM Axium Neurostimulator System to a control cohort or Retrospective* cohort trials comparing outcomes with ≥50 patients <<implanted with, treated with, using>> the subject device Abbott / SJM Axium Neurostimulator System to a control cohort or Case series (both prospective and retrospective) reporting the outcomes of ≥ 100 patients <<implanted with, treated with, using>> the subject device Abbott / SJM Axium Neurostimulator System |
| B2 | Prospective or retrospective cohort trials comparing outcomes with 30 - 49 patients <<implanted with, treated with, using>> the subject device Abbott / SJM Axium Neurostimulator System to a control cohort or Case series (both prospective and retrospective) reporting the outcomes of 30-99 patients <<implanted with, treated with, using the subject device Abbott / SJM Axium Neurostimulator System |
| C | Studies reporting outcomes of <30 patients <<implanted with, treated with, using>> the subject device Abbott / SJM Axium Neurostimulator System |
| *Studies not identified as “prospective” are considered “retrospective” |  |

### Analysis Criteria for Quality of Clinical Data

Sample text:

The appraisal criteria described in Section 10.1.1 will be used to identify data for analysis. As suggested by WG/N56FINAL:2019, the more Level 1 grades, the greater the weight of evidence provided in comparison to other data sets. As applicable: Equivalence of the devices has been established based on clinical, technical, and biological characteristics according to the <<MEDDEV 2.7/1 Rev 4 guidelines or EU MDR Annex XIV, Part A, Section 3>>.

The clinical data appraised for suitability and data contribution will be further assessed to evaluate their relative contribution towards safety and performance. The appraisal criteria described in Section above will be used to identify data for analysis. As suggested by WG/N56FINAL:2019, the more Level 1 grades, the greater the weight of evidence provided in comparison to other data sets.

In general, datasets will be weighted in the following manner:

Table 20. Level of Evidence Criteria

| Level of Evidence | Criteria |
| --- | --- |
| Level1 | Data on the <subject or equivalent device(s)>, [D1] used as intended [A1] in an applicable patient group [P1], from all study designs [T1 or T2], with high quality or minor deficiencies in acceptability / data collation [R1 or R2], outcome measures that reflect performance [O1], an appropriate follow-up period [F1], appropriate statistical analysis [S1], and a clinically significant effect [C1] were given the greatest weight in the evaluation [D1, A1, P1, T1/T2, R1 and R2, O1, F1, S1, C1]. |
| Level 2 | As it is important for manufacturers to understand off-label usage [A2, A3], use in patient groups with limited or no applicability [P2, P3], datasets pertaining to the subject devices [D1] identified during the search with high quality or minor deficiencies [R1 or R2], outcome measures that may reflect performance [O1], an appropriate follow-up period [F1], appropriate statistical analysis [S1], and a clinically significant effect [C1] were also be evaluated. To capture available data, these studies included all study design types [T1, T2, T3]. [D1/D2, A2/A3, P2/P3, T1/T2/T3, R1/R2, O1/O2, F1, S1, C1]. |
| Level 3 | Data from current knowledge/state of the art and similar competitor devices [D3] will also be analyzed, both on- and off-label [A1-A3] in all patient groups [P1-P3] with high quality or minor deficiencies [R1 or R2], all study designs [T1], outcome measures that reflect performance [O1], an appropriate follow-up period [F1], appropriate statistical analysis [S1], and a clinically significant effect [C1]. [D3, A1/A2/A3, P1/P2/-P3, T1, R1/R2, O1, F1, S1, C1]. |

Reports meeting these criteria will be critiqued and analyzed as part of the clinical evaluation.

# Analysis of Clinical Data (Benefit-Risk Analysis) (STAGE 3)

## Safety and Performance

The safety and performance measures defined in Section 2.5 for the Axium will be used in the CER to the guide the analysis of clinical data to determine if the relevant ERs are met.

Data addressing the safety and performance measures defined in Section 2.5 will be extracted from the identified sources of clinical data. The clinical data will be collectively analyzed to determine whether the Axium meet the pre-defined criteria for safety and performance.

## Risk Analysis (from Risk Management Documentation)

The Risk Management Report for the Axium will be referenced in the CER to provide a review of the risk management file and process to ensure that activities per the risk management plan have been appropriately implemented. The CER will include a summary of the risk assessment and control activities completed during device development as well as an evaluation of the overall risk acceptability and risk benefit analysis as outlined in the applicable Risk Management Reports.

The major clinical risks associated with the Axium DRG neurostimulator system include undesirable changes in stimulation that occur over time, unpleasant sensations or motor disturbances, unwanted stimulation, hemorrhage, hematoma, infection, spinal cord compression or paralysis, battery failure or battery leakage, cerebrospinal fluid (CSF) leakage, persistent pain at the electrode or IPG site, seroma at the implant site, lead migration, allergic or rejection response to implant materials, implant migration and/or local skin erosion, inflammation, headache, discomfort and inadequate pain relief. These risks will be discussed in detail in the CER.

If any new possible hazards or risks to the user or patient are identified during review of clinical data as part of the clinical evaluation process, steps will be taken to:

- Assess the risk by determining the severity of the hazard and probability of occurrence;

- Review the risk and determine the acceptable risk level;

- Control and reduce the risk to an acceptable level;

- Mitigate the risk through planned actions;

- Monitor risk reduction decisions over the product life cycle.

After all risk control measures have been implemented and verified, an evaluation shall be made to determine if the overall residual risk posed by the devices is acceptable using the criteria defined in the Risk Management Plan (RM004) and as presented in the current Risk Management Report Axium Neurostimulator System (RM006) for the devices under assessment (per Risk Management Process SOP 50-0043).

## Benefit-Risk Assessment

An overall benefit to risk analysis will be performed in the CER to evaluate whether the benefits of the Axium system outweigh the overall risks.

Any gaps will be identified in the CER if the clinical data presented and reviewed in the CER is considered insufficient to provide adequate clinical evidence in support of the anticipated conclusion that the benefits of the Axium outweigh the overall risks. A determination will be made as to the additional clinical data and data source(s) needed so that conclusions can be drawn with confidence in relation to conformity with the ERs.

## Additional Data

If the collective data from the above data sources are determined to be insufficient to make a reasonable conclusion regarding the acceptable safety and performance of the devices under consideration, the following actions/items will be performed:

- Identify any gaps or missing data

- Determine the additional clinical data and data source(s) that are needed so that conclusions can be drawn with confidence in relation to conformity with the essential requirements, including:

  - Evaluation of the safety, performance and benefit/risk profile

  - Compatibility with a high level of protection of health and safety

  - The acceptability of any undesirable side-effects

The risk of use error and the adequacy of the IFU to intended users
However, based on previous clinical evaluations of the Axium, it is not anticipated that this will be an issue.

# Summary of the Clinical Evaluation (Stage 4)

Sample text:

The clinical evidence will be critically assessed and summarized in the CER to confirm whether the safety and performance of the medical devices under assessment, when used under the conditions and for the purposes intended, as specified by the manufacturer, are in compliance with Annex I of the MDD (MDD 93/42/EEC, updated per the 2007/47/EEC M5 amendment), OR Annex 1 of the AIMDD 90/385/EEC, OR Annex I of the Medical Device Regulation (EU) 2017/745, MEDDEV 2.7.1, Revision 4 and ISO 14155:2011. Conclusions will be made in the CER with respect to the following:

- Compliance with <<ERs>> OR <<GSPRs>> (and/or Essential Principles) on clinical safety and performance of the devices under assessment when used according to their intended purpose, including the acceptability of the benefit/risk profile and of side-effects;

- Consistency and alignment between the clinical evaluation and the information materials supplied by the manufacturer (including the label, IFU, available promotional materials, including accompanying documents possibly foreseen by the manufacturer);

- Consistency between these informational materials and the current knowledge/ state of the art;

- Whether additional clinical investigations or other measures are necessary in order to generate additional data so that conclusions can be drawn regarding conformity with the <<ERs>> OR <<GSPRs>>;

- Residual risks and uncertainties or unanswered questions (including on rare complications, long term performance, safety under wide-spread use), whether these are acceptable for continued CE-mark certification, and whether they are required to be addressed during PMCF and /or PMS..

# Updating the Clinical Evaluation Plan

Sample text:

The CEP will be updated per the frequency specified in <<DOP6525/SOP 87302>> at a minimum.

# Appendix 1: Product Codes

Table 22. Product Codes

| Absolute .035 CE Part Numbers- Legal Mfg. SantaClara |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Part Name/Description | Base Part # | Sufffix (Stent Length), -XX |  |  |  |  |  |
| CE ABSOL .035 5/XX/55 | FG1010523 | -20 | -30 | -40 | -60 | -80 | -100 |
| CE ABSOL .035 5/XX/80 | FG1010531 | -20 | -30 | -40 | -60 | -80 | -100 |
| CE ABSOL .035 5/XX/135 | FG1010539 | -20 | -30 | -40 | -60 | -80 | -100 |
| CE ABSOL .035 6/XX/55 | FG1010524 | -20 | -30 | -40 | -60 | -80 | -100 |
| CE ABSOL .035 6/XX/80 | FG1010532 | -20 | -30 | -40 | -60 | -80 | -100 |
| CE ABSOL .035 6/XX/135 | FG1010540 | -20 | -30 | -40 | -60 | -80 | -100 |
| CE ABSOL .035 7/XX/55 | FG1010525 | -20 | -30 | -40 | -60 | -80 | -100 |
| CE ABSOL .035 7/XX/80 | FG1010533 | -20 | -30 | -40 | -60 | -80 | -100 |
| CE ABSOL .035 7/XX/135 | FG1010541 | -20 | -30 | -40 | -60 | -80 | -100 |

# Appendix 2: Literature Search Protocol

- Purpose

The clinical literature review protocol and associated CER process is designed to meet the requirements as described in Annex 1: Essential Requirements), AIMDD 90/385/EEC, MEDDEV 2.7.1, Revision 4 and ISO 14155.

- Scope Of Clinical Evaluation And Literature Review

This clinical literature review protocol presents methods used to identify pertinent data in the clinical literature that demonstrates the continued safety and performance of Abbott Medical’s Axium Neurostimulator System.

- Search Strategy

Database Searches

The Allied & Complementary Medicine™, BIOSIS Previews®, Embase®, EMCare®, International Pharmaceutical Abstracts, MEDLINE®, Dimensions, PubMed Central and ToxFile® scientific literature bibliographic databases will be searched. These databases were selected to provide the most comprehensive coverage of the medical and biomedical literature and to ensure that all relevant studies will be identified while minimizing selection bias.

In addition, other methods may be utilized such as “ad hoc searches” to adequately capture other relevant clinical literature. Articles that may have not appeared in the scientific literature bibliographic database results but provide pertinent information regarding Abbott Medical’s Axium Neurostimulator System safety or efficacy will be added. Articles may be identified ad hoc by reviewing the reference section of articles identified in the literature searches or through internet searches.

Search Terms

A comprehensive search strategy should include terms that cover five key areas. This formula ensures the identification of all relevant clinical safety and performance data on the subject device and/or equivalent comparator in the intended indication:

- Clinical conditions, disease, problem, indication

- Anatomical locations, intended population

- Interventions, therapies, generic device names

- Trade device names, manufacturers

- Other relevant terms

The search terms are detailed in Table 14. These preliminary search terms will be combined into technical database search terms and database queries.  Adverse events are not part of the search terms, because all clinical safety and performance outcomes will be sought in the target population irrespective of the measured or reported outcomes.

Table 14:  Preliminary Search Terms

| Search Term Categories | Specific Search Terms |
| --- | --- |
| Clinical Condition, Disease, Problem, Indication | Chronic Intractable Pain, Chronic Pain, Failed Back Surgery Syndrome, Intractable Low Back and Leg Pain, Intractable Low Back Pain, Intractable Leg Pain, Complex Regional Pain Syndrome Type I, Post Laminectomy Syndrome, Complex Regional Pain Syndrome Type II, Reflex Sympathetic Dystrophy, Intractable Neuropathic Pain, FBSS, CRPS |
| Anatomical Locations, Intended Population | Spine, Low Back, Leg, Limbs, Trunk, Dorsal Root Ganglia |
| Interventions, Therapies, Generic Device Names | Neuromodulation, Neurostimulation, Dorsal Root Ganglion Stimulation, DRG Stimulation |
| Trade Device Names, Manufacturers | SJM, St. Jude Medical, Abbott, Abbott Medical, Axium, Axium Neurostimulator System, Spinal Modulation Neurostimulator System |

Additional Search Tools

The results of the Allied & Complementary Medicine, BIOSIS Previews, Embase, EMCare, International Pharmaceutical Abstracts, MEDLINE, ToxFile, Dimensions and PubMed Central  scientific literature bibliographic database searches will be imported into an internal literature management system (ProQuest Dialog), which allows searches on the full text, abstract, keywords, and title of those imported articles by the subject device and equivalent comparator brand names and manufacturers. ProQuest Dialog is an essential tool in identifying literature on the safety and performance of subject devices and equivalent comparators, because it allows extensive searching on all relevant articles based on its ability to pull full text search results into an environment where the full text can be searched for device names and provides full text articles (as available) from search results. In addition, reference lists from review articles will also be searched for additional relevant articles.

- Screening Criteria

Titles and abstracts of references identified by the search strategy will be assessed in terms of their relevance and design according to the selection criteria outlined below in Table 15.

Table 15: Publication Inclusion and Exclusion Criteria

| Exclusion Criteria | Inclusion Criteria |  |
| --- | --- | --- |
| Technical articles (animal, biomechanical, cadaveric) |  | Articles with clinical safety or performance results for patients treated with Abbott Medical/SJM Axium Neurostimulator System for Chronic Intractable Pain |
| Review articles and meta-analyses |  |  |
| Studies where clinical data are unextractable |  |  |
| Studies where subject devices or equivalent comparators are not used or are not the main focus of the publication |  |  |
| Conference Abstract, Commentary, Letter to Editor |  |  |
| Duplicate Articles |  |  |

- Data Collection Plan

Eligibility (Screening and Selection)

The references identified by the search strategy will be assessed in terms of relevance and design according to the screening criteria above. All relevant publications will be reviewed and those considered eligible will be included.  Publications out of scope when assessed against the inclusion and exclusion criteria will be excluded with a reason for exclusion assigned.

Appraisal

The methodological quality of individual papers will be assessed based on the study size, design and the reporting of outcomes specific to the Axium Neurostimulator System using the criteria described in

Table 12.

Data Extraction

Data from the included studies will be extracted using predefined extraction tables.

Safety and Performance Data Outcomes

Once data extraction is completed, a table with all extracted data will be generated. The Data Extraction Table includes the following general elements for each included article:

  - Reference

  - Study design/evidence level

  - Study population characteristics

  - Clinical performance results

  - Clinical safety outcomes

More specifically, the clinical performance and safety outcomes will include the variables or endpoints displayed in Table 16.

Table 16: Outcomes Collected

| Variable | Variable Collected |
| --- | --- |
| Study Design | Study Objective(s) Study Design Number of Sites |
| Patient Characteristics | Diagnosis/Indications Device Type (Trial Lead, etc.) Device Brand Name Number of Patients Country Age Gender Follow-up Time |
| Performance Outcomes | 36-Item Short Form Health Survey (SF-36) EQ5D Pain Visual Analog Scale (VAS) Patient Reported Percent Pain Relief Patient Satisfaction Quality of Life Improvement McGill Pain Questionnaire Oswestry Disability Questionnaire Numeric Rating Scale Pain Catastrophizing Scale PROMIS-29 Analgesic Use Opioid Use |
| Safety Outcomes | Infection Undesirable Changes in Stimulation Pain, Redness or Numbness at the IPG Site Pain, Redness or Numbness at the Electrode Site Diminished or Loss of Pain Relief Lead Migration Battery Failure Cerebro Spinal Fluid Leakage Local Skin Erosion Seroma at Implant Site Wound Dehiscence Pulse Generator Migration Lead/Extension Breakage Reoperation Revision/Device Removal Death Other Adverse Events |

Performance outcomes and the numbers of patients will be extracted and recorded for the various study endpoint(s.) In general, outcomes will be extracted at the latest time point (i.e., longest duration of mean/median follow-up) to determine long term success. Multiple safety and performance variables will be considered for collection and analysis; however, it is not expected that all studies will have evaluated the same performance or safety outcomes.

All adverse events occurring in patients who received the subject device will be recorded. Device specific and procedure specific safety data will consist of all adverse events or complications that might or might not be related to the device, in the opinion of the investigator. When this information is not available, the data will be extracted as reported by the authors, including the numbers of patients evaluated for safety. The types and rates of safety events will be reported.

- Data Collection Analysis

Clinical outcomes from individual, included articles will be assessed for device safety and performance as well as consistency with the State of the Art (taking into account variability in treatment populations and study protocols). The clinical data will not be pooled across articles; instead ranges and rates will be summarized for individual publications.

- Type of Evaluation

This protocol documents the process for the annual update to the Abbott Medical Axium Neurostimulator System Literature Review and will cover the timeframe of 01 December 2018 through 31 December 2019 for chronic Intractable Pain.

- Device Description

The European Union Axium Neurostimulator System components which are in scope of this evaluation are outlined in Table 1.

- Background

The Clinical Literature Review, based on this protocol, will present data pertaining to the safety and performance of Abbott Medical’s Axium Neurostimulator System in the treatment of chronic, intractable pain.

- Methods

  - Define Literature Search

A comprehensive literature search will be conducted by a qualified information analyst or technical librarian to identify published articles with pertinent clinical evidence that demonstrate the continued safety and performance of Abbott Medical’s Axium Neurostimulator System. The search will be for English language articles from peer-reviewed journals that provide data from clinical trials with human subjects. The clinical trials should evaluate the efficacy of Abbott Medical’s Axium Neurostimulator System and provide information regarding adverse events or complications during the study.

The key questions the literature review intends to answer are:

- What are the clinical outcomes associated with Abbott Medical’s Axium Neurostimulator System?

- What are the complications or adverse events associated with Abbott Medical’s Axium Neurostimulator System?

A search will be performed for Chronic Intractable Pain. The search terms which will be used are outlined in Table 17.

Table 17: Literature Review Searches

| Database | Search Timeframe | Search Terms |
| --- | --- | --- |
| Chronic Intractable Pain |  |  |
| Allied & Complementary Medicine, BIOSIS Previews, Embase, EMCare, International Pharmaceutical Abstracts, MEDLINE and ToxFile | 01 December 2018 – 31 December 2019 | ti,ab,su,if,me,mq,tn(Abbott or Jude or SJM OR “Axium” OR “AxiumTM” OR “AxiumDRG” OR “AxiumDRGTM” OR “AxiumTMDRG” OR “AxiumTMDRGTM”) AND (((dorsal n/1 root n/1 gangli*) OR “DRG”) AND ( stimul* OR excit* OR puls* OR neuromodul* OR neuro-modul* OR neurostim* OR neuro-stim*)) AND pd(20181201-20191231) NOT ((child OR children OR infant$1 OR infancy or infantile OR newborn$1 OR neonat* OR baby or babies OR preschool* OR toddler$1 OR juvenile$1 OR boy$1 OR girl$1 OR p?ediatric* OR teen$1 OR teenage* OR adolescen*) OR (animal$1 OR veterinar* OR (animal* n/3 (stud* or test$1 or tested or testing or experiment*)) OR zoo$1 OR zoological OR zoology or estuar* or farm$1 OR mouse or mice or murine or rat or rats or rodent$1 or hamster$1 or chinchilla$1 or gerbil$1 or “guinea pig$1” or ferret$1 or woodchuck$1 OR bat$1 OR dog$1 or canine$1 or cat$1 or feline$1 OR rabbit$1 OR pig$1 or piglet$1 or swine or porcine or hog$1 or cow$1 or cattle or bovine OR sheep or ewe$1 or ram$1 or lamb$1 or wether$1 or ovine or goat$1 OR horse$1 or equine or (non-human-primate$1) or “nonhuman primate$1” or "nhps" or "nhp" or monkey$1 or chimp* or macaque$1 or cynomolgus or rhesus or gorilla$1 or orangutan$1 or gibbon$1 OR owl$1 or avian or bird$1 or chick$3 OR goose or geese or duck$1 OR frog$1 or toad$1 or amphibian$1 OR fish$2 or zebrafish or marine or dolphin$1 OR in-vitro or invitro OR ecolog*)) |
| Dimensions | 01 December 2018 – 31 December 2019 | (Abbott OR Jude OR SJM OR “Axium” OR “AxiumTM” OR “AxiumDRG” OR “AxiumDRGTM” OR “AxiumTMDRG” OR “AxiumTMDRGTM”) AND (“dorsal root ganglion”~2 OR “dorsal root ganglia”~2) AND (stimul* OR excit* OR puls* OR neuromodul* OR neuro-modul* OR neurostim* OR neuro-stim*) NOT (animal OR veterinar* OR “animal study”~3 OR “animal test”~3 OR “animal testing”~3 OR “animal experiment”~3 OR “animal tested”~3 OR zoo OR zoological OR zoology OR estuar* OR farm OR mouse OR mice OR murine OR rat OR rodent OR hamster OR chinchilla* OR gerbil* OR “guinea pig” OR ferret OR woodchuck OR bat OR dog OR canine OR cat OR feline OR rabbit OR pig OR piglet OR swine OR porcine OR hog OR cow OR cattle OR bovine OR sheep OR ovine OR ewe OR ram OR lamb OR wether OR goat OR horse OR equine OR “nonhuman primate” OR “non-human primate” OR “NHP” OR “NHPs” OR monkey OR chimp* OR macaque OR cynomolgus OR rhesus OR gorilla OR orangutan OR gibbon OR avian OR bird OR chick* OR goose OR geese OR owl OR duck OR amphibian* OR toad OR frog OR fish OR fishes OR zebrafish OR marine OR dolphin OR ecolog* OR “in-vitro” OR invitro) Filters: Publication Years (2018-2019), Health Category (Neurological), Publication Type (Articles) |
| PubMed Central | 01 December 2018 – 31 December 2019 | (Abbott[ti] or Jude[ti] or SJM[ti] or Abbott[ab] or Jude[ab] or SJM[ab] or Abbott[kwd] or Jude[kwd] or SJM[kwd] OR “Axium” OR “AxiumTM” OR “AxiumDRG” OR “AxiumDRGTM” OR “AxiumTMDRG” OR “AxiumTMDRGTM”) AND ("dorsal root ganglion” OR “dorsal root ganglia”) AND (stimul* OR excit* OR puls* OR neuromodul* OR neuro-modul* OR neurostim* OR neuro-stim*) AND ("2018/12/01"[PubDate] : "2019/12/31"[PubDate]) NOT (animal$1 OR veterinar* OR (animal* n/3 (stud* or test$1 or tested or testing or experiment*)) OR zoo$1 OR zoological OR zoology or estuar* or farm$1 OR mouse or mice or murine or rat or rats or rodent$1 or hamster$1 or chinchilla$1 or gerbil$1 or “guinea pig$1” or ferret$1 or woodchuck$1 OR bat$1 OR dog$1 or canine$1 or cat$1 or feline$1 OR rabbit$1 OR pig$1 or piglet$1 or swine or porcine or hog$1 or cow$1 or cattle or bovine OR sheep or ewe$1 or ram$1 or lamb$1 or wether$1 or ovine or goat$1 OR horse$1 or equine or (non-human-primate$1) or “nonhuman primate$1” or "nhps" or "nhp" or monkey$1 or chimp* or macaque$1 or cynomolgus or rhesus or gorilla$1 or orangutan$1 or gibbon$1 OR owl$1 or avian or bird$1 or chick$3 OR goose or geese or duck$1 OR frog$1 or toad$1 or amphibian$1 OR fish$2 or zebrafish or marine or dolphin$1 OR in-vitro or invitro OR ecolog*) |

  - Present Literature Review Results

The results of the clinical literature review will be presented and summarized in the Clinical Evaluation Report for the Spinal Modulation Neurostimulator System (Axium Neurostimulator System) (CR002, Rev. Q).

- Attachments

None