Clinical Evaluation Plan for Spinal Modulation Neurostimulator System (Axium Neurostimulator System)

Approval

| Role | Approver | Signature/Date (ddMMyyyy) |
| --- | --- | --- |
| Global Regulatory Operations-Clinical Evaluation (GRO-CE) /Author | Ramakrishnan Singam | See electronic signature in Windchill |
| GRO-CE PM/Director | Sangeeta Mantripragada | See electronic signature in Windchill |
| R&D/Clinical Engineering | Tj Baggett | See electronic signature in Windchill |
| Regulatory Affairs | Jordan Hanson | See electronic signature in Windchill |
| Medical Affairs | Allen Burton, MD | See electronic signature in Windchill |
| Clinical Affairs | Edward Karst | See electronic signature in Windchill |

# Revision History

| Revision | Description |
| --- | --- |
| A | Initial release |
| B | Annual update, Annual update to meet requirements of 87281, Rev C, including update of post market surveillance data, SOA and literature review to include relevant clinical data from articles published from November 1, 2017-December 15, 2018 |
| C | Annual update: Annual update to meet requirements of 87281, Rev D, including update of post market surveillance data, SOA and literature review to include relevant clinical data from articles published from December 15, 2018-December 31, 2019 |
| D | Annual update: Annual update to meet requirements of 87281, Rev E, including update of post market surveillance data, SOA and literature review to include relevant clinical data from articles published from January 01, 2020-March 31, 2021 |

# Table of Contents

1.0	Scope of Clinical Evaluation Plan (Stage 0)	9

1.1	Objective	9

1.2	Device(s) under Evaluation	10

1.3	Type of Assessment	12

1.4	Timeframe of Assessment	13

1.5	Newly Emerged Clinical Concerns	14

1.6	Summary of Prior Clinical Evaluation Activities	14

1.7	Relevant Essential Requirements	17

1.8	Applicable Regulatory References	17

2.0	Clinical Background and State of the Art (SOTA)	18

2.1	Literature Search Summary	21

2.1.1	Background	21

2.1.2	Objectives	22

2.1.3	Methods	25

3.0	Safety and Performance Measures/Objectives	29

4.0	Device Description	34

4.1	Detailed Device Description	34

4.2	Name and Address of Manufacturer	40

4.3	Principles of Operation	40

4.4	Development History	44

4.5	Regulatory History	45

4.6	Product Labeling	48

4.6.1	Intended Purpose/Use	48

4.6.2	Intended Target Groups	49

4.6.3	Indications	49

4.6.4	Contraindications	50

4.6.5	Warnings	50

4.6.6	Precautions	54

4.6.7	Potential Adverse Events	56

4.7	Intended User	57

4.8	Intended Clinical Benefits	57

4.9	Claims	58

4.10	Devices Used in Combination (if applicable)	58

4.11	Animal Tissues or Derivatives (if applicable)	58

5.0	Clinical Development Plan (Stage 0)	59

6.0	Equivalency Assessment (Stage 0)	59

7.0	Identify Pertinent Data for Clinical Evaluation (stage 1)	61

7.1	Data Sources	61

7.2	Non-Clinical Data	62

7.3	Clinical Investigations (or Justification for No Clinical Investigation)	62

7.4	Data Retrieved from Literature	62

7.4.1	Literature Search Protocol	63

7.5	Post-Market Surveillance Data	66

7.6	Post-Market Clinical Follow-Up Data	66

8.0	Appraisal of Pertinent Data for Clinical Evaluation (STAGE 2)	67

8.1.1	Appraisal Criteria	67

8.1.2	Analysis Criteria for Quality of Clinical Data	69

9.0	Analysis of Clinical Data (Benefit-Risk Analysis) (STAGE 3)	69

9.1	Safety and Performance	70

9.2	Risk Analysis (from Risk Management Documentation)	70

9.3	Benefit-Risk Assessment	71

9.4	Additional Data	72

10.0	Summary of the Clinical Evaluation (Stage 4)	72

11.0	Updating the Clinical Evaluation Plan	73

12.0	APPENDICES	74

13.0	REFERRENCES	87

# Table of Tables

Table 1. Axium Neurostimulator System (Current AIMDD Certified)	10

Table 2. Axium Products (Rationalized from EU)	11

Table 3. Timeframe of Assessment	14

Table 4. Axium Neurostimulator System CER Revision History	15

Table 5. Applicable Regulatory References	17

Table 6. PICO Terms	23

Table 7. State of the Art Search Terms	26

Table 8. Competitor Search Terms	27

Table 9. State of the Art Inclusion-Exclusion Criteria	29

Table 10. Safety and Performance Measures	31

Table 11. Name and Address of Manufacturer	40

Table 12. Axium Neurostimulator system Regulatory History	45

Table 13: Regulatory Approval History (Current AIMDD devices used in Proclaim DRG)	45

Table 14: Regulatory History of Rationalized Axium Devices	46

Table 15. Countries Where Axium Neurostimulation System Approved	48

Table 16: Device Instructions for Use Reference and All Labeling	48

Table 17. CER Data Sources	62

Table 18:  Axium Neurostimulator System Search Terms	65

Table 19. Appraisal Criteria for Suitability and Data Contribution	69

Table 20. Level of Evidence Criteria	70

Table 21. Residual Risks	72

Table 22.  Axium Models for US and Australia Markets	75

Table 23. Literature Search summary	80

Table 24:  Preliminary Search Terms	83

Table 25. Inclusion / Exclusion Criteria	85

Table 26: Outcomes Collection Plan	86

# Table of Figures

Figure 1. Illustration of TNS and INS Components (For Component Identification Only)	36

Figure 2. Epidural needle access using the loss-of-resistance technique	41

Figure 3. Initial lead insertion into the epidural space through the delivery needle	42

# Acronyms and Abbreviations

| Term | Definition |
| --- | --- |
| AE | Adverse Event |
| AIMDD | Active Implantable Medical Device Directive |
| BUDI | Basic Unique Device Identification |
| CAPAs | Corrective And Preventative Actions |
| CDP | Clinical Developmental Plan |
| CEP | Clinical Evaluation Plan |
| CER | Clinical Evaluation Report |
| CRPS | Complex Regional Pain Syndrome |
| CSF | Cerebrospinal Fluid |
| DRG | Dorsal Root Ganglion |
| EMI | Electromagnetic Interference |
| ERs | Essential Requirements |
| FBSS | Failed Back Surgery Syndrome |
| FDA | Food and Drug Administration |
| FMEA | Failure Modes and Effects Analysis |
| FSCAs | Field Safety Corrective Actions |
| ICD | Implantable Cardiac Defibrillators |
| IFU | Instruction For Use |
| IMDRF | International Medical Device Regulators Forum |
| INS | Implantable Neurostimulator |
| ISO | International Organization for Standardization |
| MAUDE | Manufacturer and User Facility Device Experience |
| MDCE WG | Medical Device Clinical Evaluation Working Group |
| MEDDEV | European Medical Device Vigilance System |
| MRI | Magnetic Resonance Imaging |
| NACC | Neurostimulation Appropriateness Consensus Committee |
| NS | Neurostimulator |
| NSAIDs | Non-Steroidal Anti-Inflammatory Drugs |
| PMCF | Post Marketing Clinical Follow-Up |
| PMS | Post Marketing Surveillance |
| RF | Radio Frequency |
| S&P | Safety and Performance |
| SAE | Serious Adverse Event |
| SCS | Spinal Cord Stimulation |
| SOP | Standard Operating Procedure |
| SOTA | State-of-the-Art |
| TENS | Transcutaneous Electrical Nerve Stimulation |
| TMS | Transcranial Magnetic Stimulation |
| TNS | Trial Neurostimulator |
| VAS | Visual Analog Scale |

# Scope of Clinical Evaluation Plan (Stage 0)

## Objective

This Clinical Evaluation Plan (CEP) summarizes the plan to continuously identify, collect, appraise and analyze relevant clinical data to provide clinical evidence in the Clinical Evaluation Report (CER) for the Abbott Axium Neurostimulator System to demonstrate continued conformity with the Essential Requirements (ER) for clinical safety and performance when used according to the manufacturer’s Instruction for Use (IFU).

The CEP discusses the plans to address the following stages of the clinical evaluation process:

Stage 0: Define the scope of the clinical evaluation

Stage 1: Identify pertinent data

Stage 2: Appraise pertinent data

Stage 3: Analyze the clinical data

Stage 4: Summarize the clinical evaluation

The clinical evaluation process is designed to meet the requirements as described in Annex 7 of the AIMDD (Active Implantable Medical Device Directive) 90/385/EEC, updated per the 2007/47/EEC M5 amendment, Annex 1: Essential Requirements, AIMDD 90/385/EEC, MEDDEV 2.7/1, Rev 4, and EN ISO 14155:2011.

## Device(s) under Evaluation

The Axium Neurostimulator System under evaluation hold current CE-Mark and are commercially marketed in the EU as well as rationalized Axium products. Models registered and/or commercially marketed in other geographies such as US and Australia are included in Appendix 1.

- This plan applies to the following products in Table 1 and Table 2. The models currently certified under AIMDD do not have BUDI determinations. Hence, the Basic UDI (BUDI) is not applicable (N/A) currently.

- Slim Tip Leads: lead model numbers ending with an “A” suffix

- Ball Tip Leads: lead model numbers with no suffix (no longer manufactured)

Table 1 lists the model numbers that are included in the scope of the clinical evaluation.

Table 1. Axium Neurostimulator System (Current AIMDD Certified)

| Basic UDI-DI | Model Number | Device Name | EU Device Classification / Rule |
| --- | --- | --- | --- |
| IPG Programmers and Accessories |  |  |  |
| NA | MN20700 | Clinical Programmer | AIMD, Class III; Rule 8 |
| NA | MN20600-02 | Patient Programmer | AIMD, Class III; Rule 8 |
| NA | MN23300 | Auxiliary Magnet Kit | AIMD, Class III; Rule 8 |
| IPG Chargers and Accessories |  |  |  |
| NA | MN23400 | Clinical Programmer Charger Kit | AIMD, Class III; Rule 8 |
| NA | MN23400-U | Patient Programmer Charger | AIMD, Class III; Rule 8 |
| NA | MN23500 | Clinical Programmer Carrying Case | AIMD, Class III; Rule 8 |
| NA | MN23500-S | Patient Programmer Carrying Case | AIMD, Class III; Rule 8 |
| Leads & Lead Accessories |  |  |  |
| NA | MN20450-50A, MN20450-90A | Implantable SlimTip Lead, 50cm and 90cm* | AIMD, Class III; Rule 8 |
| NA | MN20550-50 | Lead Extension Kit* | AIMD, Class III; Rule 8 |
| NA | MN22150 | 22cm Small Curve Delivery Sheath Kit* | AIMD, Class III; Rule 8 |
| NA | MN23650 | 22cm Big Curve Delivery Sheath Kit* | AIMD, Class III; Rule 8 |
| NA | MN24000 | Curved Needle* | AIMD, Class III; Rule 8 |
| NA | MN22050 | DRG Lead Accessories Kit* | AIMD, Class III; Rule 8 |
| NA | MN21350 | DRG Connector Cable Kit* | AIMD, Class III; Rule 8 |

*These models were exclusively used with the Axium Neuromodulation System until 2016, when Proclaim DRG was initially released for market distribution. Between 2016 and 2018, their usage have been with Proclaim DRG and Axium Systems. Axium devices manufactured by Spinal Modulation ceased being manufactured in December 2017. These models are covered in the Proclaim DRG CER 90222654 where their usage after 2016 is surveilled.

Table 2. Axium Products (Rationalized from EU)

| Basic UDI-DI | Model Number** | Device Name | EU Device Classification / Rule |
| --- | --- | --- | --- |
| IPG and EPG |  |  |  |
| NA | MN20200 | Implantable Neurostimulator | AIMD, Class III; Rule 8 |
| NA | MN20100 | Trial Neurostimulator | AIMD, Class III; Rule 8 |
| Leads & Lead Accessories |  |  |  |
| NA | MN20450-50 MN20450-90 | Implantable Lead (ball tip), 50cm and 90cm | AIMD, Class III; Rule 8 |
| NA | MN20350-50 MN20350-90 | Trial Lead (ball tip), 50cm and 90cm | AIMD, Class III; Rule 8 |
| NA | MN20350-50A MN20350-90A | Trial SlimTip Lead Kit, 50cm and 90cm | AIMD, Class III; Rule 8 |
| NA | MN21900 MN22200 | Tunneling Tool Kit 30cm/51cm | AIMD, Class III; Rule 8 |
| NA | MN23850 | 22cm Axium Small Curve Delivery Sheath Kit | AIMD, Class III; Rule 8 |
| NA | MN23950 | 22cm Axium Big Curve Delivery Sheath Kit | AIMD, Class III; Rule 8 |
| Programmer& Programmer Accessories |  |  |  |
| NA | MN20600 | Patient Programmer | AIMD, Class III; Rule 8 |

** Axium products which are not marketed or sold in EU.

## Type of Assessment

The clinical evaluation is routine (regularly scheduled) annual update to assess the clinical safety and performance of the device to demonstrate conformity with the relevant ERs 1, 2 and 5 per AIMDD. The CER will include an evaluation of the safety, performance, risk/benefit profile, and acceptability of undesirable side-effects as identified through a review of the following:

- Description of device(s) and any design features of the device(s), or any indications or target populations that require specific attention.  This also covers any design features that pose special performance or safety concerns, the intended purpose and application of the device(s) (e.g. target treatment group and disease, proposed warnings, contraindications, precautions, and method of application) and the specific claims made by the manufacturer about the clinical performance and clinical safety of the subject device(s);

- Data from risk management documents. This clinical evaluation will assess the significance of any residual clinical risks that remain after design risk mitigation strategies have been employed.

- The current knowledge/state of the art in the corresponding medical field, such as applicable standards and guidance documents, information relating to the medical condition managed with the device and other treatment alternatives available to the target population.

- Relevant clinical data on safety and performance collected through Abbott sponsored pre- and/or post-market clinical studies, peer-reviewed published literature, and post-market surveillance data. Both favorable and unfavorable data shall be taken into account

## Timeframe of Assessment

The timeframe and period of the current assessment is noted below in Table 3.

Table 3. Timeframe of Assessment

| Category | Time Frame | Date Range | Justification |
| --- | --- | --- | --- |
| Literature Search | 1 Year | 01-January-2020 -. 31-March-2021 | This is an annual update hence a 1-year timeframe is selected for evaluation. The date-range for this evaluation is selected to ensure there are no gaps in data analyzed relative to the previous (year 2019) evaluation #CR002. Rev R |
| PMS (Post Marketing Surveillance) Data | 1 Year | Through 31-March-2021 (Previous search period through December 31, 2019) | One-year timeframe is used to enable appropriate data for trending and complaint analysis. The date-range for this evaluation is selected to ensure there are no gaps in data analyzed relative to the previous (year 2019) evaluation #CR002. Rev R |

## Newly Emerged Clinical Concerns

The CER update is provided for the purposes of annual update and is not the result of newly emerged clinical concerns.

## Summary of Prior Clinical Evaluation Activities

This CEP is for the update to the existing CER on Axium Neurostimulator System, and will be Clinical Evaluation Report  CR002 Rev S. CER revision history is shown below in Table 4.

Table 4. Axium Neurostimulator System CER Revision History

| Revision | Description |
| --- | --- |
| A | Initial Release. |
| B | Before: A revision history was provided, but no history overview was included. After: A History Overview has been added to the Appendix to provide ‘Before’ and ‘After’ history for each change. Page numbers were updated. |
| C | Before: No indication language for Nevro was included. Appendix F did not contain a column explaining similarities/differences between devices. No post-market surveillance plan or clinical study reports were included. Need for an additional literature review and conclusion by an impartial expert. After: Appendix F has been updated to be more comprehensive. Indications of Nevro were added and a more detailed comparison of the indications for different SCS companies was added. A post-market surveillance plan and clinical study report summarizing the current clinical trial data were included. Added letters of support from Dr. Timothy Deer and Dr. Liong Liem. |
| D | Add signature sheet to front of report, remove discussion of off-label use in patient registry patients, add units of measure for Charge and Charge Density in Appendix F. |
| E | Change post-market follow-up duration from 3 months to 30 days. |
| F | Add post-market study performance, safety, and effectiveness data; update the Instructions for Use; update the risk analysis data; update format to MEDDEV 2.7.1 r 3, Appendix E; Remove Appendices I, J and K; add new appendices I, J and K; Update Appendix C with DRG publications from Spinal Modulation Inc. |
| G | Update literature search added dorsal column stimulation/DCS to search terms in the Spinal Cord Stimulation section, updated appendices with new lit search data, updated references with new references, added new CV for lit search personnel. Added new US indication language. Updated to current component configurations and materials. Updated complaints table (thru Q3 2013). |
| H | Update literature search, updated appendices with new literature search data tables, updated SCS/PNS/DRG reference section with new references. Deleted “Ongoing EU Studies” appendix and moved into text (Table 15). Added CV of lit search personnel. Deleted Appendix G (of the previous draft): History Overview. Renumbered appendices. Added SlimTip performance data (Appendix H), DRG bibliography through 04/07/2015 (Appendix K) and vigilance reports (Appendix L). |
| J | Update literature search, updated appendices with new literature search data tables, updated SCS/PNS/DRG reference section with new references, added new CV for literature search personnel. Updated complaints table (Appendix J) and vigilance reports (Appendix L). Added sales numbers for the product (Appendix M). |
| K | Updated literature, updated appendices with new literature search data tables, updated SCS/PNS/DRG reference sections with new references, added new CV for literature search personnel. Updated complaints table (Appendix J) and vigilance reports (Appendix L). Added sales numbers for the product (Appendix M). Updated PMCF reference. Updated Table 6. Updated cardiac implantable device contraindication. |
| L | Added information for US models to Appendices J, L, and M. |
| M | Generalized “Intended Therapeutic Indications” and “Contraindications” statements; updated status of post-market studies in Table 6. |
| N | Updated CER to be compliant with MEDDEV 2.7.1, Rev. 4. |
| P | 2019 Annual Clinical evaluation Update; updated literature from November 1, 2017-December 15, 2018 and post market surveillance data including complaints, FSCA, and CAPAs through November 2018. |
| Q | 2020 Annual Clinical evaluation update, update literature from 01 December 2018 – 31 December 2019 and post market surveillance data including complaints, FSCA and CAPAs through December 31, 2019 |
| R | Updated CER: Clinical Evaluation Report for the Spinal Modulation Neurostimulator System (Axium Neurostimulator System) Change: Annual Update. No Rev Q was written. |
| S | Annual update: Annual update to meet requirements of 87281, Rev E, including update of post market surveillance data, SOA and literature review to include relevant clinical data from articles published from January 01, 2020-March 31, 2021 |

## Relevant Essential Requirements

This CER will address and provide support for conformance to the requirements on Safety and Performance as per ERs 1, 2, and 5 as per AIMDD. Additionally, this clinical evaluation will also support conformance to other essential performance using relevant and specified clinical outcome parameters to evaluate the intended clinical benefits to patients, as presented in  Section 4.8.

The Essential Requirements Checklist(s) for each product in scope of this evaluation satisfy will be reviewed and evaluated for continued conformance to the Essential Requirements per AIMDD (90506719).

## Applicable Regulatory References

Table 5. Applicable Regulatory References

| Standards Organization and Reference Number | Title |
| --- | --- |
| MDD 93/42/EEC | Council Directive 93/42/EEC of 14 June 1993 Concerning Medical Devices |
| AIMDD 90/385/EEC | Council Directive 90/385/EEC; Annex 7, updated per the 2007/47/EEC M5 amendment, concerning Active Implantable Medical Devices |
| MEDDEV 2.7/1 Rev 4 | European Commission Guidelines on Medical Devices: Clinical Evaluation: A Guide for Manufacturers and Notified Bodies under Directives 93/42/EEC and 90/385/EEC |
| IMDRF MDCE WG/N56 Final:2019 | International Medical device Regulator’s Forum Medical Device Clinical Evaluation Working Group’s Guidance on Clinical Evaluation |
| ISO 13485:2016 | Medical devices -- Quality Management Systems -- Requirements for Regulatory Purposes |
| EN ISO 14155:2011 | Clinical Investigation of Medical Devices for Human Subjects Good Clinical Practice |
| EN ISO 14971:2012 | Medical devices - Application of Risk Management to Medical Devices |
| EN ISO 10993-1 | Biological evaluation of medical devices – Part 1: Evaluation and testing within a risk management process |

# Clinical Background and State of the Art

The CER will include a description of the clinical background, treatment options and current knowledge /State-of-the-Art (SOTA) on the treatment of management of chronic intractable pain.

Relevant guidelines on the disease state/medical conditions/patient population/ treated by the Axium Neurostimulator system and clinical literature on the Standard-of-Care/alternative treatment options/competitor device on the Axium Neurostimulator system itself will be identified through the literature search.

The CER will include a discussion on the Standard of Care/ Relevant benchmark guidelines and relevant clinical evidence on the subject device and competitors/comparators/ alternative treatment options. Data on safety and performance of key/relevant comparators/competitors will be used to characterize the placement of the Axium Neurostimulator system in the state of the art. Key competitors considered for this purpose will include Nevro, Senza® and Medtronic, Recharging Restore™. The alternate treatment option categories for chronic intractable pain such as Surgical/Non-Surgical treatment will be summarized in CER.

## Literature Search Summary

A comprehensive, objective/non-biased literature searches will be conducted to identify relevant State-of-the-Art literature. Recent reviews and meta-analysis/systematic reviews will be identified through these searches may also be leveraged as data-leads to identify relevant literature through back-referencing.

### Background

Axium Neurostimulator is intended to use for the treatment of management of chronic intractable pain by delivering the low-intensity electrical impulses to nerve structures in line with SOTA. The literature review will provide data on current interventions (Includes therapies, diagnostic measures, measures for the management of diseases or medical conditions) for the intended patient population (State-of-the-Art) in order to give input to the assessments of acceptable benefit/risk profiles, what is currently considered as providing a high level of protection of health and safety and what are considered acceptable side-effects.

Broad literature searches (e.g., chronic back pain, chronic leg pain, chronic pain, chronic pain management) which focused on systematic reviews and meta-analyses will be conducted in addition to internet searches. Findings will be reviewed and included if they provided the latest treatment options and outcomes for the treatment and management of patients with chronic intractable pain. More targeted searches will also be conducted including Failed Back Surgery Syndrome (FBSS), Complex Regional Pain Syndrome (CRPS) and Dorsal Root Ganglion (DRG) Neurostimulation. In addition, publications from the systematic literature searches will be reviewed to identify treatment options and interventions.

### Objectives

The objective of the systematic literature review is to provide a comprehensive analysis of the available published clinical literature with the objective to asses and analyze the body of evidence relating to the SOTA.

The SOTA literature search strategy on chronic intractable pain and DRG stimulation was developed based on information relevant to the intended patient population, interventions, comparative interventions and outcomes (PICO; Table 6).

Table 6. PICO Terms

| P population | I intervention | C compare | O outcomes |
| --- | --- | --- | --- |
| Patients with chronic intractable pain, including, but not limited to, Failed Back Surgery Syndrome, Intractable Low Back and Leg Pain, Intractable Low Back Pain, Intractable Leg Pain, Complex Regional Pain Syndrome Type I, Post Laminectomy Syndrome, Complex Regional Pain Syndrome Type II, Reflex Sympathetic Dystrophy, Intractable Neuropathic Pain, FBSS, CRPS | Education & Exercise Cognitive Behavior Therapy Acupuncture Physical Therapy TMS (Transcranial Magnetic Stimulation) TENS (Transcutaneous Electrical Nerve Stimulation) NSAIDs (Non-Steroidal Anti-Inflammatory Drugs) Muscle Relaxants Antispastics Antidepressants Gabapentinoids Tramadol Opioids Epidural Steroid Injections Facet Interventions Nerve Block Ablation Sacroilliac Injections Adhesiolysis Decompression Fusion Spinal Cord Stimulators (SCS) Dorsal Root Ganglion (DRG) Stimulators Intrathecal Drug Pumps | Nevro, Senza Medtronic, Recharging Restore™ | Pain relief, reduction in pain, patient satisfaction, activity level, reduction in pain medication use, complication, quality of life |

### Methods

In addition to previously identified SOTA articles from literature searches and reviews conducted for past CERs, a systematic and comprehensive literature search will be performed to identify new information on current SOTA practice guidance and expert consensus statements relevant to the treatment and management of chronic intractable pain, including DRG stimulation. The SOTA review of published literature will include practice guidelines, consensus documents, and competitive device articles to obtain information on clinical background and current knowledge in the relevant medical fields and disciplines relating to chronic intractable pain, and information relevant to treatment / management options including DRG stimulation. Information and data from this SOTA will be reviewed to evaluate safety, effectiveness and acceptable side effects of treatment.

The database used for literature retrieval is Dialog Databases which covered Allied & Complementary Medicine™, Embase®, EMCare®, International Pharmaceutical Abstracts, MEDLINE®, ToxFile®. The literature articles will be selected based on the Inclusion/Exclusion criteria as mentioned Table 9. The literature screening process and Appraisal plan will be performed based on Literature search protocol provided in Appendix 2.

Table 7. State of the Art Search Terms

| Set# | Database | Search Timeframe | Searched for SOTA |
| --- | --- | --- | --- |
| S3 – Combined SOA and Guidelines | Dialog Databases: Allied & Complementary Medicine™, Embase®, EMCare®, International Pharmaceutical Abstracts, MEDLINE®, ToxFile® | 01 January 2020 – 31 March 2021 | S1 OR S2 |
| S2 - Guidelines |  |  | ti,ab((spine$1 OR spinal) n/3 (stimulat$5 OR neurostimulat$5 OR neuromodulat$5 OR excit* OR puls*)) AND ti(guideline$1 OR consensus OR statement$1 OR “task force$1” OR recommend* OR (position$1 n/3 (statement$1 OR paper$1 OR document$1 OR societ$4 OR chapter)) OR ((societ$4 OR organization$1) n/5 (guideline$1 OR guidance OR consensus OR statement$1 OR paper$1 OR recommend*)) OR (clinic$5 n/3 (practice$1 OR advisor$4)) OR NACC OR "Neuromodulat$5 Appropriate$5 Consensus Committee") AND pd(20200101-20210331) and la(English) |
| S1 – SOA |  |  | ti(causalgia OR (peripher$3 n/3 neuropath*) OR ((chronic$4 OR intractab$6 OR continu$6 OR persist$5 OR constant$3 OR sustain$4 OR visceral$2 OR burning OR back OR discogenic OR disc$1) n/3 (pain$1 OR painful$3)) OR ((complex$3 OR compound$5 OR complicat$4) n/3 region$4 n/3 pain$1) OR CRPS$1 OR (fail$4 n/3 back n/3 surger$3) OR FBSS) AND (ti((DRG OR (dorsal n/1 root n/1 gangli$3)) n/3 (stimulat$5 OR neurostimulat$5 OR neuromodulat$5 OR excit$6 OR puls$6)) OR ti(acupuncture OR ((epidural$1 OR sacroili$2 OR sacrum-ili$2 OR facet$1 OR joint$1) n/3 (implant$5 OR inject$5 OR administr$6 OR administer$3 OR steroid$2)) OR ((neural OR nerv$5) n/3 (lesion$4 OR strip$5 OR block$4 OR denervat$4)) OR neurotom$3 OR rhizotom$3 OR neurolys?s OR adhesioly$4 OR opioid OR opioids OR opiate OR opiates OR NSAID OR NSAIDs OR tramadol OR gabapentin OR gabapentinoid OR ((muscle OR muscular$2) n/3 relax$4) OR antispastic$1 OR anti-spastic$1 OR antispasmodic$1 OR anti-spasmodic$1 OR antidepressant$1 OR anti-depressant$1 OR CBT OR (cognitive n/1 behavior$2 n/3 therap$3) OR ((intrathecal OR ((spine OR spinal OR cerebrospinal) n/1 fluid$1) OR CSF) n/3 (pump OR deliver$3)) OR (disc$1 n/3 (compress$5 OR decompress$5)) OR ((spine OR spinal) n/3 (fusion$1 OR fused OR fusing)) OR neuroplast$3 OR TMS OR (((transcranial$2 n/3 magnet$6) OR (transcutaneous$2 n/3 nerv$5) OR ((spine OR spinal) n/3 cord$1)) n/3 (stimulat$5 OR neurostimulat$5 OR neuromodulat$5 OR excit$6 OR puls$6)) OR TENS OR SCS OR exercis$4 OR (physical$2 n/3 (activ$6 OR therap$3)))) AND (ti,ab(RCTs OR (random* n/2 control$5)) OR ti,ab,dtype(((control$5 OR random*) n/5 (stud$5 OR trial$3)) OR (systematic$5 n/1 review$5) OR meta-analys?s OR metaanalys?s OR metanalys?s)) AND pd(20191001-20200930) and la(English) NOT (ti,ab(pediatric$4 OR paediatric$4 OR child OR children OR infant$1 OR infancy OR infantile OR newborn$1 OR neonat* OR baby OR babies OR babe$1 OR preschool* OR toddler$1 OR boy$1 OR girl$1 OR teen$1 OR teenage$3 OR adolescen$4 OR youth$1 OR juvenile$1 OR veterinar$5 OR zoo$1 OR zoological OR zoology OR estuar* OR farm$1 OR mouse OR mice OR murine OR rat OR rats OR rodent$1 OR hamster$1 OR chinchilla$1 OR gerbil$1 OR “guinea pig$1” OR ferret$1 OR woodchuck$1 OR bat$1 OR dog$1 OR canine$1 OR pup$1 OR puppy OR puppies OR cat OR cats OR feline$1 OR rabbit$1 OR hare$1 OR bunny OR bunnies OR pig$1 OR piglet$1 OR gilt$1 OR sow$1 OR boar OR boars OR hog$1 OR swine OR porcine OR cow$1 OR cattle OR calf OR calves OR bovine OR sheep OR ewe$1 OR ram$1 OR lamb$1 OR wether$1 OR ovine OR (ovis p/0 aries) OR goat$1 OR horse$1 OR mare$1 OR foal$1 OR equine OR equestr* OR (non-human-primate$1) OR “nonhuman primate$1” OR "nhps" OR "nhp" OR monkey$1 OR chimp* OR macaque$1 OR cynomolgus OR rhesus OR gorilla$1 OR orangutan$1 OR gibbon$1 OR in-vitro OR invitro OR ecolog* OR enviro* OR species OR spp) OR dtype(patent$5 OR chapter$1 OR correct$5 OR erratum$1 OR editor$5 OR comment$5 OR letter$1 OR response$1 OR note$1 OR survey$1 OR (case$1 n/3 (analys?s OR report$1 OR stud$5 OR series)) OR “conference abstract$1” OR “meeting abstract$1”)) |
| #2 | Pubmed | 01 January 2020 – 31 March 2021 | < ((causalgia OR (peripher* neuropath*) OR (chronic intractab* OR visceral OR burning OR back OR discogenic OR disc) AND (pain OR painful OR CRPS OR (fail back surger*) OR FBSS) AND (DRG OR (dorsal root gangli*) OR stimulat* OR neurostimulat* OR neuromodulate*) ) AND (("2020/01/01"[Date - Publication] : "2021/03/31"[Date - Publication])) Filters: Humans, English > |

Table 8. Competitor Search Terms

| Database | Search Timeframe | Search Terms |
| --- | --- | --- |
| #1 (Dialog) | 01 January 2020 – 31 March 2021 | (Senza* OR OR Senza OR Senza TM* OR Senza?C$2 AND Nevro) AND ti,ab,su,if,me,mq,tn,co(brain OR neuro* OR AND pd(20200101-20210331) and la(English) NOT (ti,ab(veterinar* OR zoo$1 OR zoological OR zoology OR estuar* OR farm$1 OR mouse OR mice OR murine OR rat OR rats OR rodent$1 OR hamster$1 OR chinchilla$1 OR gerbil$1 OR “guinea pig$1” OR ferret$1 OR woodchuck$1 OR bat$1 OR dog$1 OR canine$1 OR pup$1 OR puppy OR puppies OR cat OR cats OR feline$1 OR rabbit$1 OR hare$1 OR bunny OR bunnies OR pig$1 OR piglet$1 OR gilt$1 OR sow$1 OR boar$1 OR hog$1 OR swine OR porcine OR cow$1 OR cattle OR calf OR calves OR bovine OR sheep OR ewe$1 OR ram$1 OR lamb$1 OR wether$1 OR ovine OR (ovis p/0 aries) OR goat$1 OR horse$1 OR mare$1 OR foal$1 OR equine OR equestr* OR (non-human-primate$1) OR “nonhuman primate$1” OR "nhps" OR "nhp" OR monkey$1 OR chimp* OR macaque$1 OR cynomolgus OR rhesus OR gorilla$1 OR orangutan$1 OR gibbon$1 OR owl$1 OR avian OR bird$1 OR chick$3 OR goose OR geese OR duck$1 OR frog$1 OR toad$1 OR amphibian$1 OR fish$2 OR zebrafish OR marine OR aquatic OR dolphin$1 OR in-vitro OR invitro OR model* OR ecolog* OR enviro* OR parasit* OR entomolog* OR insect* OR larva* OR species OR spp OR alga OR algae OR daphnia OR plant* OR botan* OR horticultur* OR tree$1) OR dtype(patent* OR “conference abstract$1” OR “meeting abstract$1” OR “conference poster$1” OR “meeting poster$1” OR chapter$1 OR correct* OR erratum* OR editor* OR comment* OR letter$1 OR response$1 OR note$1 OR survey$1)) Limiters Applied: See Exclusion/Inclusion Criteria |
| #2 (Pubmed) |  | (((Nevro Senza) OR (Nevro Senza®) OR (Nevro® Senza®) OR (Senza) OR (Senza®)) AND (Medtronic Recharging Restore™) OR (Recharging Restore™) OR (Medtronic, Recharging Restore™) AND ((neurostimulation) OR (neurostimulat*) OR (SCS) OR (spinal cord stimulation) OR (spinal cord stim*) OR (neuromodulation) OR (neuromodulat*)) ) AND (("2020/01/01"[Date - Publication] : "2021/03/31"[Date - Publication])) Filters: Humans, English |

Table 9. State of the Art Inclusion-Exclusion Criteria

| Exclusion Criteria | Inclusion Criteria |
| --- | --- |
| (E1) Does not meet SOTA inclusion criteria | (I1) Articles informing on the state of the art regarding dorsal root and spinal cord stimulation, the relevant patient population, or treatment alternatives (I2) Articles presenting new/state of the art information regarding DRG Stimulation devices including device safety and/or performance relevant to the state of the art, when used as intended similar to subject devices I3) Competitive - Article reports benchmark clinical study data of a competitive device in >30 patients - Article reports clinical outcomes of safety and performance of a competitive device when used as intended Clinical outcomes reported are specific to the competitor device (i.e., results not pooled with other devices) |

# Safety and Performance Measures/Objectives

Table 10 summarizes the safety and performance objectives, measures and acceptance criteria that have been established for the Axium Neurostimulator system.

Table 10. Safety and Performance Measures

| Objective | Measure/Acceptance Criteria |
| --- | --- |
| Safety The Axium Neurostimulator system is safe when used as intended to treat patients with chronic intractable pain that has been unresponsive to other methods. | Measure Incidence of serious adverse events related to the Axium Neurostimulator system. Acceptance Criteria for Serious Adverse Event (SAE)/ Adverse Event (AE) Acceptable safety criteria are established based on a severe adverse event rate ranging from 0 to 13.1% reported in literature and clinical study results (with n ≥ 30 patients implanted)1 |
| Performance The Axium Neurostimulator system is intended to provide pain relief in patients experiencing chronic intractable pain. | Measures -Patients achieve ≥ 30% reduction in pain from baseline -Patients achieve a ≥ 2-point increase in PROMIS-29 physical function2 Acceptance Criteria The incidence of pain relief from the Axium Neurostimulator system is expected to be consistent with that found in the clinical practice guidelines. Pain relief ≥ 30% has been noted as a clinically significant difference3 Pain relief is change on a pain intensity scale from baseline to follow-up. Improvement in physical function is expected to be consistent with literature pointing to a clinically significant difference of 2 points on the PROMIS physical functioning scale4 |

Note: Acceptance criteria for the safety measure utilized the following SAE rates from clinical study data where number of implanted patients ≥30. ACCURATE study results were published in Deer et al.1

The safety and performance objectives and measures will be identified based on a review of clinical practice guidelines, state of the art, competitive and subject device literature consistent with the subject device indications for use. The Axium Neurostimulation System’s safety measure and acceptance criteria will be determined to be consistent with the SAE and AE rates identified in the clinical practice guidelines and literature relating to implantation of Axium Neurostimulator into the epidural space of the brain.

Similarly, the performance measure acceptance criteria relating to successful deliver of low-intensity electrical impulses to nerve structures will be measured based on SOTA, as well as subject and competitive device literature related to successful implantation of Axium Neurostimulator in patients with Chronic pain.

# Device Description

## Detailed Device Description

The Axium Neurostimulator System is composed of two subsystems as pictured in Figure 1. The Axium Neurostimulator System is a fully implantable medical device that stimulates neural tissue to produce analgesia for the treatment of chronic pain.

Figure 1. Illustration of TNS and INS Components (For Component Identification Only)

The following is a description of each of the components of the Axium Neurostimulator System:

- Trial Neurostimulator (TNS)

The external TNS device connects to the Trial Lead(s) or Lead Extensions and is worn by the patient for up to 30 days during the trial period. The TNS device has a belt clip for the patient’s convenience.

- Implantable Neurostimulator (INS)

The Axium INS is a non-rechargeable, 4-channel electronic device. It uses microelectronic circuitry, powered by a hermetically sealed battery, to generate a pulsed waveform to stimulate neural tissue. The electronic circuitry and battery are housed in a hermetically sealed titanium case.

Each neurostimulator has a unique internal identifier that allows the physician to identify the type of device through an X-ray. The radiopaque identifier inside the case allows identification of both the device manufacturer and model number using standard x-ray equipment. For the Axium Neurostimulator, the code is SM001, which identifies Spinal Modulation as the manufacturer.

The INS is packaged in a sealed inner tray within a sealed outer tray.

- Implant Leads/Trial Leads/Lead Extension

The Lead Kits contain the Leads and the individual delivery devices that are required for their placement.

Implant / Trial Leads: The Leads are designed for percutaneous introduction into the body using a special needle and a set of custom delivery tools provided in their respective kits. Each Lead is fitted with four cylindrical electrodes spaced at equal intervals, which are intended to provide stimulation at the target DRG. The temporarily placed Trial Leads are intended for use for up to 30 days.

Lead Extension: The Lead Extension consists of a silicone port header that accepts the Axium Trial Lead and Implant Lead. It is intended to extend the length of the lead and provide a connection between the lead and the Connector Cable, or the lead and the Implantable Neurostimulator System. The Lead Extension is intended for chronic implantation as a component of the Axium Neurostimulator System.

- Connector Cable Kit

The Connector Cable connects the Leads or Lead Extension to the external TNS. The Connector Cable is packaged separately from the Lead and Lead Extension Kit. The Connector Cable includes a connector and 2 extension cables for use as needed.

The Lead Identifiers are small bent wires that can be inserted into the proximal stylet lumen in the leads and are individually colored to identify the leads.

- Lead Accessories

Small / Big Curve Delivery Sheaths are intended to allow passage of the lead percutaneously into the epidural space. The labeled length of the sheath is the distance from the hub to the pre-shaped tip and the length of the curve at the tip is approximately 2 mm for the Small Curve and approximately 8mm for the Big Curve.

Axium Small / Big Curve Delivery Sheaths are also intended to allow passage of leads percutaneously into the epidural space. Axium sheaths are internally reinforced with thin stainless-steel braiding.

Complex Curve and Straight Stylets assist in steering and positioning the lead within the epidural space. These are included within the 50cm and 90cm SlimTip trial lead kits.

The 14G Delivery Needle is intended to access the epidural space, providing a conduit for lead, guidewire and delivery sheath placement. It is available as a straight needle or a curved needle. The Axium Curved Needle is available in a separate package and is uniquely identified by the colored hub. The 14G Delivery Needle is only available in the Implant Lead, Trial Lead, and Lead Accessories Kits.

The Guidewire is optional and is intended to verify that the needle is in the epidural space after using a loss of resistance technique.

Soft Tissue Anchors are intended to anchor the Lead in the soft tissue or on the skin surface proximal to the distal contacts of the Lead.  The soft tissues anchors are packaged with the leads and included in the lead accessories kit.

- Implantation Tools

The tunneling tool is used to provide a conduit for the Trial Lead, Implant Lead, or Lead Extension to the INS or away from the midline of the spine. It is packaged with 2 exchangeable tips: a blunt pencil tip and a sharp trocar tip. A straw is slid over the tunneling tool and when the steel handle is removed, the straw provides the conduit for tunneling.

The INS Sizer is approximately the same size as the INS and allows the physician to properly size the INS pocket. The INS sizer is packaged with the Axium tunneling tool kits (MN21900, MN22200).

The port plugs are used to fill unused ports in the INS. The port plugs are packaged with the INS, but spare port plugs are also packaged with the Tunneling Tool Kits for the convenience of the physician.

- Additional Accessories

Sterile Magnet Sleeve: The magnet is placed in the sterile sleeve to allow it to be used during the implantation of the INS. The sterile magnet sleeve is packaged with the Axium tunneling tool kits (MN21900, MN22200).

Medical Alert Card identifies the patient as a user of the Neurostimulation System. The Medical Alert card is packaged with the INS and Patient Programmer.

Programmer Charger is to be used with the Clinical or Patient Programmers to charge the battery or allow use of the Programmers while plugged into standard electrical outlets.

Programmer Carrying Case protects the Programmers from liquid.

Auxiliary Magnet allows the user to turn the NS off or activates RF to allow the user to communicate with the NS.

Hex Key allows the user to release a set screw in the INS header or Lead Extension header that has been unscrewed too far. The Hex Key is not to be used to tighten the set screw against the lead. The hex key is packaged with the Axium tunneling tool kits (MN21900, MN22200).

- Clinical Programmer and Patient Programmer

The Clinical Programmer is used to program the stimulation parameters for both the TNS and the INS. The instructions for programming the TNS and INS devices are the same. The Clinical Programmer is used by the physician or clinical staff.

The Patient Programmer allows the patient to adjust the stimulation settings of the TNS and INS devices within limits preset by the clinician. The Patient Programmer also allows the patient to turn stimulation off, if necessary.

## Name and Address of Manufacturer

The legal manufacturer and manufacturing site are listed in Table 11.

Table 11. Name and Address of Manufacturer

| Device name | Legal Manufacturer | Manufacturing Site |
| --- | --- | --- |
| Axium Neurostimulator system | Abbott Medical | 6901 Preston Road Plano, TX 75024, USA |

## Principles of Operation

Identical to SCS, DRG stimulators are composed of electrical contacts (leads) and an implanted pulse generator (IPG). Electrical pulses are delivered via a lead that is percutaneously implanted at the DRG (in contrast to the dorsal columns in SCS). The implanted lead is powered by a neurostimulation generator, which generates the electrical pulses and is surgically implanted under the skin. Stimulation of the DRG modifies the patient experience of neuropathic pain by dampening the firing of the aberrantly firing DRG neurons. The operating principle of the electrode component of percutaneous leads is to convert the charges carried by electrons in the lead conductor wires into charges carried by the ions in the electrolytes of the biological tissue. The ultimate dose of electrical stimulation is the resulting combination of pulse width, pulse frequency, and current amplitude. These parameters can be modulated using the clinician programmer, pulse amplitude and on/off times are adjusted in order to maximize the pain suppression for an individual patient. The patient can turn the stimulator on or off and may vary the stimulation parameters within pre-set limits, as required, using a hand-held remote control (patient programmer).

Pulses are delivered via a lead that is implanted surgically or percutaneously, at the DRG. The implanted DRG lead is powered by a neurostimulator device, which generates the electrical pulses and is surgically implanted under the skin. Stimulation of the DRG modifies the patient experience of neuropathic pain by dampening the firing of the aberrantly firing painful DRG neurons; it can replace painful sensations with a tingling sensation defined as paresthesia, but the analgesia is largely paresthesia-free. The electrode must be carefully positioned by a physician so that the DRG sub serving the area where pain is experienced is stimulated.

The Axium Neurostimulator System involves a two-step process for its use.  The first step is a patient trial period, where leads are percutaneously implanted into the epidural space and the patient is fitted with the externalized Trial Neurostimulator for up to 30 days. The goal of this period is to first establish whether the patient will benefit from the therapy before fully implanting the neurostimulator. The trial period typically lasts 3-30 days. If the patient does not experience sufficient pain relief in 3 days, the physician may opt to continue the trial period for no longer than 30 days.

The procedure of inserting and placing the Spinal Modulation leads is similar to that of other spinal cord stimulation devices. The delivery accessories packaged in the Trial Lead Kit or Implant Lead Kit will aid the physician with lead placement. The patient will undergo a minimally invasive surgical procedure with local anesthetic and light sedation (monitored anesthesia care) during which the leads will be inserted and steered towards the appropriate intraspinal location under fluoroscopic guidance. The physician first accesses the epidural space percutaneously using a 14 Gauge Delivery Needle (Figure 2).  This is equivalent to the method used in currently approved spinal cord stimulation systems. The lead is inserted into the sheath and both are fed through the Delivery Needle (Figure 3) and steered towards the lateral epidural space. Once the lead is in position, the delivery tools (sheath, stylet, and needle) are removed, leaving the lead in place.

Figure 2. Epidural needle access using the loss-of-resistance technique

Figure 3. Initial lead insertion into the epidural space through the delivery needle

Up to four leads may be implanted at different locations in the lateral epidural space. The number of leads placed, and their delivery locations will depend upon the anatomic location of the patient’s specific pain symptoms and will be determined by the physician.  The physician may choose to implant either the Trial Leads or the Implant Leads in the trial phase as they are identical in materials, design, and construction. This is standard practice in current spinal cord stimulation therapy. Trial leads can be connected directly to the TNS and the implant leads are connected to a lead extension inside the body. The lead extension exits the body to be connected to the TNS. This ensure that the implant lead remains sterile during the trial phase.

The patient will be awake and conversant during the procedure to provide feedback to the physician. Following placement, each lead will be connected to the Trial Neurostimulator using the Connector Cable.  Under direction of the physician, trained clinical staff will use the Clinical Programmer to communicate wirelessly with the TNS and adjust the stimulation settings (pulse amplitude, pulse width, frequency, and lead electrode polarity). To ensure appropriate lead positioning and stimulation settings, the patient will provide feedback to the physician on the effect of stimulation (paresthesia) on their pain. Once the final lead position has been determined, leads will be held in place using the Soft Tissue Anchors. Trial Leads will be sutured to the skin. Implant Leads will be sutured to the soft tissue near the spinous processes and a lead extension will be attached and tunneled away from the needle insertion site. The leads exit the skin and the external lead exit point will be protected per usual standard of care. The patient then returns to daily activity to evaluate the effects of stimulation.  Each patient will use a hand-held Patient Programmer to adjust stimulation amplitude as needed throughout the trial period.

The TNS is worn by the patient for up to 30 days.  At the end of the trial period, the physician and patient will determine if the therapeutic effects warrant implantation of the INS. As is standard practice in spinal cord stimulation therapy, patients achieving sufficient pain relief with the Trial Neurostimulator will have the option to get the neurostimulator implanted. Following the trial period, the TNS and Connector Cable will be disconnected from the leads and the Trial Leads will be removed from the patient’s epidural space.  If the physician had previously placed the Implant Leads during the trial, they will remain in place. If the trial stimulation was not successful, all leads will be removed, and the patient will not receive the implantable device.

To proceed with the Implantable Phase, the patient will be scheduled for an implantation procedure similar to the trial procedure. If Implant Leads were used during the trial period, these leads will be tunneled to a pocket in the abdomen or upper buttocks. If Trial Leads were used during the trial period, the patient will undergo a second lead implantation during which Implant Leads will be placed using the same percutaneous techniques as in the Trial Lead implant procedure. Once the leads are placed, the physician will confirm proper positioning by connecting them to a TNS. The patient will again give feedback to the physician on the effect of stimulation on their pain and, if needed, the physician can revise lead position. The leads will then be secured to the tissue, using the Soft Tissue Anchors, and tunneled, using the Tunneling Tool, through the tissue under the skin to the INS pocket. The INS will be placed in this pocket and attached to either the Implant Leads themselves, or, if needed, a Lead Extension that connects the lead to the INS. The pocket and incision site will be sutured closed.

The subject will return home with an INS and up to four leads implanted in the epidural space. Due to the design of the device, this therapy is fully reversible and can be removed through a similar, minor procedure. Clinical staff will use the Clinical Programmer to program the stimulation settings of the INS and the patient will be given a Patient Programmer that they can use to adjust stimulation settings within pre-set limits determined by the physician.

## Development History

The Axium Neurostimulator System for the treatment of chronic intractable pain was approved for commercial distribution in Europe in November 2011. In addition, the Axium Neurostimulator System for the treatment of chronic, intractable pain of the trunk and/or limbs was approved for commercial distribution in Australia in 2013. In 2014, communications and manual updates were sent to the European and Australian regulators and physicians with regard to the appropriate technique for lead removal. The Axium Neurostimulator System was approved for commercial distribution in the United States in February 2016.  The device has not been withdrawn from marketing for any reason related to its safety or effectiveness.

The Axium Neurostimulator System was originally approved with Spinal Modulation (located in Menlo Park, California) as the legal manufacturer and manufacturing facility.  In May 2015, Spinal Modulation became a wholly owned subsidiary of St. Jude Medical. St. Jude Medical developed and received CE marking approval for the next-generation DRG system (Proclaim DRG) in September 2016.  The Proclaim DRG system utilizes some of the same DRG leads and accessories as the Axium Neurostimulator System.

In order to consolidate facilities, the DRG leads, accessories and Axium programmers were transferred from Spinal Modulation to St. Jude Medical’s facility in Plano, Texas as the legal manufacturer and manufacturing facility. The Axium Implantable Neurostimulator, Trial Neurostimulator and other Axium-specific devices were not transferred to Plano as they are phasing out due to the Proclaim DRG system. The Spinal Modulation facility ceased manufacturing and permanently closed its doors in December 2017. All CE marked devices manufactured by Spinal Modulation were placed on the EU market for distribution and the Spinal Modulation CE certificate was cancelled in March 2018. Tables 4 and 5 outline the CE marking history and current certificates for the devices in the Axium Neurostimulator System.

## Regulatory History

The CE Mark approval and certification timelines for the Axium Neurostimulator system are shown in Table 12.

Table 12. Axium Neurostimulator system Regulatory History

| Device Name | Model number | GMDN Code | Date of First CE Mark | Certificate Number (SJM / Abbott) |
| --- | --- | --- | --- | --- |
| Clinical Programmer | MN27000 | 61588 | 16 Nov 2011 | CE 639587 CE 706798 |
| Patient Programmer | MN20600-02 | 61588 | 04 Mar 2014 |  |
| Auxiliary Magnet Kit | MN23300 | 57996 | 16 Nov 2011 |  |
| Clinical Programmer Charger Kit | MN23400 | 17115 | 16 Nov 2011 |  |
| Patient Programmer Charger | MN23400-U | 17115 | 04 Mar 2014 |  |
| Clinical Programmer Carrying Case | MN23500 | 12500 | 16 Nov 2011 |  |
| Patient Programmer Carrying Case | MN23500-S | 12500 | 04 Mar 2014 |  |

NOTE: Several DRG products are rebranded to Abbott Medical.  For the devices indicated, we have mirror CE certificates – one for Abbott Medical, one for St. Jude Medical to facilitate the rebranding transition.

Table 13: Regulatory Approval History (Current AIMDD devices used in Proclaim DRG)

| Device Name | Model number | GMDN Code | Date of First CE Mark | Certificate Number (SJM / Abbott) |
| --- | --- | --- | --- | --- |
| Implantable SlimTip Lead, 50cm and 90cm* | MN20450-50A, MN20450-90A | 36007 | 11 Dec 2014 | CE 639589 CE 691529 |
| Lead Extension Kit* | MN20550-50 | 36007 | 16 Nov 2011 |  |
| 22cm Small Curve Delivery Sheath Kit* | MN22150 | 44063 | 16 Nov 2011 |  |
| 22cm Axium Big Curve Delivery Sheath Kit* | MN23650 | 44063 | 16 Nov 2011 |  |
| Curved Needle* | MN24000 | 58293 | 04 Oct 2013 |  |
| Lead Accessory Kit* | MN22050 | 44063 | 16 Nov 2011 |  |
| Connector Cable Kit | MN21350 | 47143 | 16 Nov 2011 |  |

*These models are all Axium models that prior to 2016 were exclusively used with the Axium Neuromodulation System. When Proclaim DRG was initially released for market distribution in 2016, their usage may have been with either Proclaim DRG or Axium Systems. Axium devices manufactured by Spinal Modulation ceased being manufactured in December 2017.

NOTE: Several DRG products are rebranded to Abbott Medical.  For the devices indicated, we have mirror CE certificates – one for Abbott Medical, one for St. Jude Medical to facilitate the rebranding transition.

Table 14: Regulatory History of Rationalized Axium Devices

| Device Name | Model Number | GMDN Code | Date of First CE Mark | Current Certificate |
| --- | --- | --- | --- | --- |
| Trial Neurostimulator | MN20100 | 36007 | 16 Nov 2011 | No, Certificate expired 05MAR2018 |
| Implantable Neurostimulator | MN20200 | 36007 | 16 Nov 2011 | No, Certificate expired 05 Mar 2018 |
| Trial SlimTip Lead Kit, 50cm and 90cm | MN20350-50A, MN20350-90A | 36007 | 11 Dec 2014 | No, Certificate expired 05 Mar 2018 |
| Implantable Lead (Ball Tip), 50cm and 90cm | MN20450-50, MN20450-90, MN20350-50, MN20350-90 | 36007 | 16 Nov 2011 | No, Certificate expired 05 Mar 2018 |
| Tunneling Tool Kit, 30cm/51cm | MN21900, MN22200 | 44063 | 16 Nov 2011 | No, Certificate expired 05 Mar 2018 |
| Patient Programmer | MN20600 | 61393 | 16 Nov 2011 | No, Certificate expired 05 Mar 2018 |
| 22cm Axium Small Curve Delivery Sheath Kit | MN23850 | 44063 | 04 Sep 2013 | No, Certificate expired 05 Mar 2018 |
| 22cm Axium Big Curve Delivery Sheath Kit | MN23950 | 44063 | 04 Sep 2013 | No, Certificate expired 05 Mar 2018 |

The Axium Neurostimulator System is currently approved to be marketed in the following countries in Table 15.  Please note that this table reflects the approval status for the Axium Neurostimulator System prior to the certificate cancellation in March 2018. All CE marked inventory was placed on the EU market prior to cancellation of the certificate and the remaining inventory is being consumed. Axium products marketed and sold only in the US and Australia are listed in Appendix 1.

Table 15. Countries Where Axium Neurostimulation System Approved

| EUROPE | GLOBAL |
| --- | --- |
| Austria, Belgium, Denmark, Finland, France, Germany, Ireland, Italy, Luxembourg, Netherlands, Norway, Spain, Sweden, Switzerland, United Kingdom | Australia, Hungary, Kuwait*, New Zealand, Qatar, South Africa, United States |

*Kuwait will accept Axium products shipped to Ministry of Health hospitals

## Product Labeling

The information on device labeling relevant to this section is provided below.

Table 16: Device Instructions for Use Reference and All Labeling

| DEVICE NAME | IFU DOCUMENT # |
| --- | --- |
| Axium Patient Programmer | Patient Programmer User Manual: ARTMT600010299 |
| Axium Clinical Programmer | Clinical Programmer User Manual: ARTMT600010365A |
| Patient Users Guide | Patient Brochure: ARTMT600010362A |
| Lead and Extension Kits | Clinician's Manual: ARTEN600021271 B |
| Axium Magnet | Axium Magnet IFU: LP1082 |

### Intended Purpose/Use

This Axium Neurostimulator System is designed to deliver low-intensity electrical impulses to nerve structures. The system is intended to be used with leads and associated extensions that are compatible with the system.

Clinical Programmer: It is intended to be used by the clinical investigator or a St. Jude Medical representative to query and program the Neurostimulator (NS), to retrieve data from the NS and to allow for adjustment of the patient’s therapy.

Patient Programmer: It is used to program the patient’s Trial or Implantable Stimulator.

Trial leads: It is intended to be used with Abbott Medical neurostimulation systems for DRG stimulation by connecting to a compatible trial stimulator, either directly or with a compatible lead extension during a trial period that shall not exceed 30 days.

Implant leads: It is intended to be used with Abbott Medical neurostimulation systems for DRG stimulation by connecting to a compatible pulse generator, either directly or with a compatible lead extension.

Lead extensions: It is intended to add length to a lead used with Abbott Medical neurostimulation systems for DRG stimulation. Lead extensions can be temporarily externalized to join a lead to a trial stimulator, or permanently internalized to join a lead to a compatible pulse generator.

Axium Auxiliary Magnet: It may be used to initiate communication between the neurostimulator and the Patient or Clinical Programmer.

### Intended Target Groups

The Axium Neurostimulator System is intended to treat patients with chronic intractable pain that has been unresponsive to other methods.

### Indications

The Axium Neurostimulator System is indicated for the management of chronic intractable pain.

### Contraindications

Patients contraindicated for the Axium Neurostimulator System are those who:

- Have an active implantable medical device including but not limited to cardiac pacemakers and cardiac defibrillators

- Are unable to operate the system

- Are poor surgical risks

- Are pregnant

- Are under the age of 18

### Warnings

The following general warnings apply to the Axium Neurostimulator System is provided in IFU mentioned in Table 16.

External Defibrillators – Safety for use of external defibrillator discharges on a patient receiving neurostimulation has not been established.

Magnetic Resonance Imaging – The patient should be advised to not undergo any elective magnetic resonance imaging (MRI) with the system in place. Use of MRI in the vicinity of the lead(s) may result in forceful dislodgment of the lead(s), or damage to the Neurostimulator. If a voltage is induced through the lead, it may cause uncomfortable (“jolting” or “shocking”) levels of stimulation or injury to the patient. MRI may cause heating at the lead tip and unintended stimulation could result in tissue damage.

CT Scans – If the patient requires a CT scan, all stimulation should be turned OFF prior to the procedure. If stimulation is not turned off, the patient may experience a momentary increase in stimulation, which may be uncomfortable. After the procedure, stimulation should be turned back on and the system should be checked for proper function.

Ultrasonic Scanning – Ultrasonic equipment may cause mechanical damage to the lead if used directly over the site.

Electrosurgery Devices – Electrosurgery devices should not be used in close proximity to implanted lead(s). Contact between an active lead and the electrosurgical pencil can cause direct stimulation of the contacted nerve and can cause severe injury to the patient. Electrosurgery devices may also damage the lead and cause a loss of stimulation.

Pediatric Use – Use of neurostimulation has not been approved specifically for patients under 18 years of age.

Pregnancy – Safety and effectiveness of neurostimulation for use during pregnancy and nursing have not been established.

Non-Emergency Procedures – The patient must be advised that they must not have non-emergency procedures while they are undergoing trial stimulation.

Emergency Procedures – The patient should be instructed to designate a representative (family member or close friend) to notify any emergency medical personnel of their neurostimulator implant, if emergency care is required. Each patient will be provided with a Medical Alert Card to carry with them that will inform emergency medical personnel of the patient’s implant. The patient should be advised to use caution when undergoing any procedure that could include RF (Radio Frequency) or microwave ablation, defibrillation or cardioversion.

Routine Medical Procedures – The patient should be instructed to not undergo dental procedures, diathermy, electrolysis, diagnostic ultrasound, static field therapeutic magnets, diagnostic X-ray, and high output ultrasonic lithotripsy. These procedures may provide interference that can affect TNS or INS device operation or use or damage components of the system that may cause patient harm. If the patient with an INS or TNS device is subsequently given any medical treatment in which an electrical current is passed through his/her body from an external source, either the device should first be deactivated, or care should be taken to monitor the functioning of the neurostimulator during the initial stages of treatment.

Diathermy Therapy – Do not use short-wave diathermy, microwave diathermy, or therapeutic ultrasound diathermy (all now referred to as diathermy) on patients implanted with a neurostimulation system. Energy from diathermy can be transferred through the implanted system and cause tissue damage at the location of the implanted electrodes, resulting in severe injury or death. Diathermy is further prohibited because it may also damage the neurostimulation system components. This damage could result in loss of therapy, requiring additional surgery for system removal and replacement. Injury or damage can occur during diathermy treatment whether the neurostimulation system is turned on or off. All patients are advised to inform their health care professionals that they should not be exposed to diathermy treatment.

Cardiac Pacemakers and Implantable Cardiac Defibrillators (ICD) – It is possible that the device may affect the operation of other implantable devices such as pacemakers or implantable cardiac defibrillators. The physician should be aware of any other implantable devices the patient may have or is scheduled to receive.

Explosive or Flammable Gases – Do not use the patient programmer or clinical programmer to communicate with the INS or TNS in an environment where explosive or flammable gas fumes or vapors are present. The operation of the programmer could cause them to ignite, causing severe burns, injury, or death.

Case Damage – If the INS case is pierced or ruptured, an explosion can occur from the battery chemicals, which can lead to severe burns or even death.

Device Components – The use of non-Spinal Modulation components with the system may result in damage to the system and increased risk of harm to the patient.

Component Disposal – Dispose of leads and neurostimulators per local requirements. Do not crush, puncture, or burn the Neurostimulator because explosion or fire may result.

Exposure to Fluids – Exposure of the external TNS or the Connector Cable to water, body fluids, saline, or cleaning agents can cause corrosion and affect stimulation. If this occurs, dry all components thoroughly prior to lead connection. Do not immerse the external TNS or Connector Cable in fluids.

Device Handling – The patient must be instructed to not remove their Trial Lead(s) or Connector Cable. Manipulation of the components may result in an undesired outcome, such as the patient developing an infection, getting undesirable stimulation, or accidentally turning their stimulation off. The patient must be instructed to not rub or exert pressure on the implantable neurostimulator through the skin as this may cause: lead dislodgement leading to stimulation at the implant site, device inversion leading to the inability to communicate with the device, or skin erosion that can lead to another surgical procedure or possible infection.

Equipment Operation – During stimulation, patients should be advised to not operate potentially dangerous equipment, such as power tools, automobiles, or other motor vehicles. They should not climb ladders or participate in other activities where postural change or an abrupt movement could alter the perception of stimulation intensity and cause patients to fall or lose control of equipment or vehicles or injure others.

Patient Activity – Patients should be advised to limit their activities to low or moderate levels during their trial stimulation period and the first six weeks of implantation of the INS. Failure to do so may result in migration of the leads causing loss of stimulation therapy, muscle stimulation or painful stimulation thereby requiring reoperation to reposition.

Theft Detectors and Metal Screening Devices – Certain types of antitheft devices, such as those used at entrances/exits of department stores, libraries, and other public establishments, and/or airport security screening devices may affect stimulation. It is possible that patients who are implanted with non-adjacent multiple leads and/or patients who are sensitive to low stimulation thresholds may experience a momentary increase in their perceived stimulation, which has been described by some patients as uncomfortable or jolting. It is recommended that patients use caution when approaching such a device and that they request assistance to bypass the device. If they must proceed through the device, patients should turn off the neurostimulator and proceed with caution, being sure to move through the detector quickly.

Restricted Areas – The patient should be warned to seek medical guidance before entering environments which could adversely affect the operation of the implanted device, including areas protected by a warning notice preventing entry by patients fitted with a pacemaker.

Electromagnetic Interference (EMI) – Certain commercial equipment such as resistance welders and induction furnaces, certain communication equipment such as microwave transmitters, linear power amplifiers, and high voltage power lines may generate sufficient EMI to interfere with the INS and TNS System resulting in under stimulation, overstimulation, or difficulty in device communications. Certain medical devices may also cause EMI interference and make it difficult to communicate with the INS or TNS; such devices include bone growth stimulators, TENS, and implanted cardiac devices.

Lead Movement – The patient should be instructed to avoid excessive bending, twisting, and stretching, entering hyperbaric chambers above 2.0 atmosphere, and operating the neurostimulator while driving or lifting objects of over 2 kg (5 lbs.) for a minimum of 6 weeks after implantation. These activities may cause lead movement, which can result in under stimulation or overstimulation for the patient. Excessive lead migration may lead to the need for reoperation to replace the leads.

Therapeutic Radiation – Therapeutic radiation may damage the electronic circuitry of an implanted neurostimulation system, although no testing has been performed and no definite information on radiation effects is available. Sources of therapeutic radiation include x-rays, cobalt machines, and linear accelerators. If radiation therapy is required, the area over the implanted INS should be shielded with lead.

### Precautions

The following precautions apply to the Axium Neurostimulator is provided in IFU mentioned in Table 16.

Physician Training – Implanting physicians should be experienced in the diagnosis and treatment of chronic pain syndromes and have undergone surgical and device implantation training.

Patient Selection – It is extremely important to select patients appropriately for neurostimulation and that thorough psychiatric screening be performed. Patients should not be dependent on drugs and should be able to operate the spinal cord stimulator system.

Infection – It is important to follow proper infection control procedures. Infections related to system implantation might require that the device be explanted.

Implantation of Two Systems – If two systems are implanted, ensure that at least 15 cm (6 in) separates the implanted INSs to minimize the possibility of interference during programming.

Implantation of Multiple Leads – If multiple leads are implanted, leads and extensions should be routed in close proximity. Nonadjacent leads can possibly create a conduit for stray electromagnetic energy that could cause the patient unwanted stimulation.

High Stimulation Outputs – Stimulation at high outputs may cause unpleasant sensations or motor disturbances or may render the patient incapable of controlling the patient programmer. If unpleasant sensations occur, the device should be turned off immediately.

Stimulation Parameters – Patients should be cautioned that stimulation parameters must be determined under the supervision of a physician and that they should not adjust stimulation parameters within prescribed programs unless ordered to do so by a physician.

Cellular Phones – The effect of cellular phones on neurostimulation systems is unknown; patients should avoid placing cellular phones directly over the system.

Trial Leads – Trial Leads are intended for use for up to 30 days. Use of these devices must be performed in accordance with the instructions provided in this manual.

Over programming – Excessive communication with the device can shorten the life of the INS. The patient should be warned to communicate with the device only when necessary.

TNS Device Care – The patient must be instructed to not spill fluids on, to wash or otherwise get their TNS device wet. The patient must not shower or bathe with it (sponge baths are acceptable as long as the TNS device does not get wet). The patient must be instructed not to drop or mishandle their TNS device. Physical damage to the unit may impair its function. The patient must be instructed to not open the TNS case.

Lead and Cable Care – The patient must be instructed to not remove their lead(s) or Connector Cable during the trial period. Manipulation of the components may result in the patient developing an infection, getting undesirable stimulation, or accidentally turning their stimulation off.

TNS Device Failure – Device failure, although unlikely, is possible due to random component failure. If the TNS device stops working, the patient should contact their physician.

TNS Device Disposal – The patient is to be instructed that they must return their TNS device and Patient Programmer to their physician after the trial period. The patient must be instructed to not discard or burn their TNS device. Fire may cause the internal battery to explode.

TNS Battery Replacement – It is unlikely that the battery will need replacement in the short time that the patient has their TNS device. However, if the TNS device does not function the patient must not try to open the TNS case. The internal battery must be replaced by Abbott Medical personnel only. The patient should be advised to contact their physician during regular business hours.

Material Sensitivity – Hypersensitivity (redness in the area of skin contact) can happen if the patient has an allergic reaction to the materials. If this occurs, the patient should be instructed to contact their physician during regular business hours.

### Potential Adverse Events

In addition to those risks commonly associated with surgery, the following risks are also associated with use of a neurostimulation system:

- Undesirable changes in stimulation, which may be related to cellular changes in tissue around the electrodes, changes in electrode position, loose electrical connections, or lead failure or breakage

- Stimulation in unwanted places

- Lead migration, causing changes in stimulation or reduced pain relief

- Epidural hemorrhage, hematoma, infection, spinal cord compression, or paralysis from placement of a lead in the epidural space

- Cerebrospinal fluid (CSF) leakage

- Tissue or nerve damage

- Paralysis, weakness, clumsiness, numbness, sensory loss, or pain below the level of the implant

- Pain or bleeding where the needle was inserted

- Persistent pain at the electrode site

- Seroma (mass or swelling) at the implant site

- Headache

- Allergic or rejection response to device or implant materials

- Implant migration or skin erosion around the implant

- Pain from a no-injurious stimulus to the skin or an exaggerated sense of pain

Additional risks, as a result of the placement and stimulation of the lead in the area of the dorsal root ganglion, include pain from setting the stimulation parameters too high. This may occur once the lead is in place and is connected to the neurostimulator and activated. The neurostimulator is controlled by a trained operator and the starting point for the stimulation will be set to the lowest available settings. Additionally, all patients will be awake and conversant during the procedure to minimize the impact.

## Intended User

Implanting physicians should be experienced in the diagnosis and treatment of chronic pain syndromes and have undergone surgical and device implantation training.

## Intended Clinical Benefits

Previous clinical evaluations show that the Axium DRG neurostimulation therapy benefits patients with chronic intractable pain who have no other options and that subjects are satisfied with the treatment. The safety profile has remained similar to the results from other studies and should be considered as an acceptable risk given these patients have failed multiple therapies before their neurostimulator implant.  Factors such as quality of life and improved sleep are also positively affected by DRG stimulation. The safety profile remains similar to the results from the pre-market study and should be considered as an acceptable risk given these patients failed multiple therapies before their neurostimulator implant. The current clinical evaluation will assess the Axium system for continued benefit and risk analysis and results will be reported in the CER.

## Claims

No specific clinical claims supporting safety and/or performance have been identified for the Axium Neurostimulator System.

## Devices Used in Combination (if applicable)

There are no other devices or products are used in combination with Axium Neurostimulator System.

## Animal Tissues or Derivatives (if applicable)

There are no components of animal origin in the Axium Neurostimulator System. Any chemicals, such as mold release agents considered as tallow derivatives have been processed under conditions at least as rigorous as those specified in ISO 22442-1, Annex C, Section C.5, are considered exempt from Commission Regulation 722/2012 of 08 August 2012 (Article 1 Section 4(a)). These are considered acceptable for Transmissible Spongiform Encephalopathies risk due to their processing conditions and not their sourcing, irrespective of the geographical origin and the nature of the tissues from which tallow derivatives are derived.

# Clinical Development Plan (Stage 0)

The Clinical Developmental Plan (CDP) is an MDR requirement and the subject device Axium Neurostimulator System remain certified under the AIMDD hence this CDP is not applicable.

# Equivalency Assessment (Stage 0)

Equivalence to other devices will not be claimed for the Axium Neurostimulator System.

# Identify Pertinent Data for Clinical Evaluation (stage 1)

Data for the clinical evaluation of Axium Neurostimulator system consists of critically-evaluated scientific literature for Axium Neurostimulator system and current knowledge/state of the art; and post-market clinical experience data including patient-effect complaints and device complaints for Axium Neurostimulator system will be included, as appropriate.

## Data Sources

Comprehensive methods will be utilized to identify clinical and non-clinical data relevant for the clinical evaluation of Axium Neurostimulator system. These include non-clinical biocompatibility evaluation and bench-top testing data, data from critically evaluated scientific literature, clinical investigations, current knowledge / SOTA; and post-market clinical experience data from patient-effect and device complaints, external vigilance data, post-market surveillance plans, and post-market clinical follow-up. Table 17 lists each data source and the sections detailing their description.

Table 17. CER Data Sources

| Data Source | Section (Detailed Description) |
| --- | --- |
| Non-clinical data (Bench-Top Testing Data, Biocompatibility Testing, Pre-clinical Studies) | Section 7.3 |
| Clinical Investigation Data | Section 7.4 |
| Literature | Section 7.5 |
| Post Marketing Clinical Follow-Up (PMCF) Data | Section 9.0 |
| Customer Complaint Data | Section 8.0 |

## Non-Clinical Data

No new verification and validation studies were performed on any models in scope of this clinical evaluation. All products described in the CER have successfully passed formal verification and validation and evaluated at the initial CE submission.

There are no new pre-clinical studies during this clinical evaluation. All pre-clinical studies for the products in scope of this evaluation are covered in the existing CER CR002.

## Clinical Investigations (or Justification for No Clinical Investigation)

Pre-Market Clinical Investigations: There are no new pre-market clinical investigations during this clinical evaluation. The completed pre-market clinical investigation studies are covered in CER CR002.

## Data Retrieved from Literature

Literature searching is used to identify data not held by the manufacturer that are needed for the clinical evaluation.

Literature searching identifies potential sources of clinical data for establishing clinical data relevant to the device under evaluation. . Published clinical safety and performance data will be retrieved for the Axium Neurostimulator system. A summary of the literature search protocol is provided below. Refer to Appendix 2 for the full literature search protocol.

### Literature Search Protocol

#### Background

A systematic literature search will be conducted to identify clinical data relevant to the safety and performance of the Axium Neurostimulator system. The review of published literature will describe the clinical background and current knowledge in the relevant medical fields and disciplines relating to the Axium. Data on the safety and performance, risk/benefit profile, and acceptability of undesirable side-effects of Axium Neurostimulator system will be analyzed. These data will be used to inform on what is currently considered to provide a high level of protection of health and safety, and what are considered acceptable side effects. Special attention will be given to the recognition and identification of potential new risks, including hazards due to substances and technologies.

#### Objective

The purpose of the systematic literature review is to provide a comprehensive analysis of the available published clinical literature to asses and analyze the body of evidence relating to clinical safety and performance data of Abbott Medical’s Axium Neurostimulator System.

#### Methods

A systematic literature review will be conducted for safety and performance of Axium Neurostimulator system using the Allied & Complementary Medicine™, BIOSIS Previews®, Embase®, EMCare®, International Pharmaceutical Abstracts, MEDLINE®, ToxFile®, Dimensions and PubMed Central® scientific literature bibliographic databases.

A summary of the literature search, screening, inclusion/exclusion criteria, Data appraisal plan and extraction methods are summarized in Appendix 2: Literature search protocol.

Table 18:  Axium Neurostimulator System Search Terms

| Database | Search Timeframe | Search Terms |
| --- | --- | --- |
| Allied & Complementary Medicine, BIOSIS Previews, Embase, EMCare, International Pharmaceutical Abstracts, MEDLINE and ToxFile | 01 January 2020 – 31 March 2021 | ti,ab,su,if,me,mq,tn(Abbott or Jude or SJM OR “Axium” OR “AxiumTM” OR “AxiumDRG” OR “AxiumDRGTM” OR “AxiumTMDRG” OR “AxiumTMDRGTM”) AND (((dorsal n/1 root n/1 gangli*) OR “DRG”) AND ( stimul* OR excit* OR puls* OR neuromodul* OR neuro-modul* OR neurostim* OR neuro-stim*)) AND pd(20200101-20210331) NOT ((child OR children OR infant$1 OR infancy or infantile OR newborn$1 OR neonat* OR baby or babies OR preschool* OR toddler$1 OR juvenile$1 OR boy$1 OR girl$1 OR p?ediatric* OR teen$1 OR teenage* OR adolescen*) OR (animal$1 OR veterinar* OR (animal* n/3 (stud* or test$1 or tested or testing or experiment*)) OR zoo$1 OR zoological OR zoology or estuar* or farm$1 OR mouse or mice or murine or rat or rats or rodent$1 or hamster$1 or chinchilla$1 or gerbil$1 or “guinea pig$1” or ferret$1 or woodchuck$1 OR bat$1 OR dog$1 or canine$1 or cat$1 or feline$1 OR rabbit$1 OR pig$1 or piglet$1 or swine or porcine or hog$1 or cow$1 or cattle or bovine OR sheep or ewe$1 or ram$1 or lamb$1 or wether$1 or ovine or goat$1 OR horse$1 or equine or (non-human-primate$1) or “nonhuman primate$1” or "nhps" or "nhp" or monkey$1 or chimp* or macaque$1 or cynomolgus or rhesus or gorilla$1 or orangutan$1 or gibbon$1 OR owl$1 or avian or bird$1 or chick$3 OR goose or geese or duck$1 OR frog$1 or toad$1 or amphibian$1 OR fish$2 or zebrafish or marine or dolphin$1 OR in-vitro or invitro OR ecolog*)) |
| Dimensions | 01 January 2020 – 31 March 2021 | (Abbott OR Jude OR SJM OR “Axium” OR “AxiumTM” OR “AxiumDRG” OR “AxiumDRGTM” OR “AxiumTMDRG” OR “AxiumTMDRGTM”) AND (“dorsal root ganglion”~2 OR “dorsal root ganglia”~2) AND (stimul* OR excit* OR puls* OR neuromodul* OR neuro-modul* OR neurostim* OR neuro-stim*) NOT (animal OR veterinar* OR “animal study”~3 OR “animal test”~3 OR “animal testing”~3 OR “animal experiment”~3 OR “animal tested”~3 OR zoo OR zoological OR zoology OR estuar* OR farm OR mouse OR mice OR murine OR rat OR rodent OR hamster OR chinchilla* OR gerbil* OR “guinea pig” OR ferret OR woodchuck OR bat OR dog OR canine OR cat OR feline OR rabbit OR pig OR piglet OR swine OR porcine OR hog OR cow OR cattle OR bovine OR sheep OR ovine OR ewe OR ram OR lamb OR wether OR goat OR horse OR equine OR “nonhuman primate” OR “non-human primate” OR “NHP” OR “NHPs” OR monkey OR chimp* OR macaque OR cynomolgus OR rhesus OR gorilla OR orangutan OR gibbon OR avian OR bird OR chick* OR goose OR geese OR owl OR duck OR amphibian* OR toad OR frog OR fish OR fishes OR zebrafish OR marine OR dolphin OR ecolog* OR “in-vitro” OR invitro) Filters: Publication Years (2020-2021), Health Category (Neurological), Publication Type (Articles) |
| PubMed Central | 01 January 2020 – 31 March 2021 | (Abbott[ti] or Jude[ti] or SJM[ti] or Abbott[ab] or Jude[ab] or SJM[ab] or Abbott[kwd] or Jude[kwd] or SJM[kwd] OR “Axium” OR “AxiumTM” OR “AxiumDRG” OR “AxiumDRGTM” OR “AxiumTMDRG” OR “AxiumTMDRGTM”) AND ("dorsal root ganglion” OR “dorsal root ganglia”) AND (stimul* OR excit* OR puls* OR neuromodul* OR neuro-modul* OR neurostim* OR neuro-stim*) AND ("2020/01/01"[PubDate] : "2021/03/31"[PubDate]) NOT (animal$1 OR veterinar* OR (animal* n/3 (stud* or test$1 or tested or testing or experiment*)) OR zoo$1 OR zoological OR zoology or estuar* or farm$1 OR mouse or mice or murine or rat or rats or rodent$1 or hamster$1 or chinchilla$1 or gerbil$1 or “guinea pig$1” or ferret$1 or woodchuck$1 OR bat$1 OR dog$1 or canine$1 or cat$1 or feline$1 OR rabbit$1 OR pig$1 or piglet$1 or swine or porcine or hog$1 or cow$1 or cattle or bovine OR sheep or ewe$1 or ram$1 or lamb$1 or wether$1 or ovine or goat$1 OR horse$1 or equine or (non-human-primate$1) or “nonhuman primate$1” or "nhps" or "nhp" or monkey$1 or chimp* or macaque$1 or cynomolgus or rhesus or gorilla$1 or orangutan$1 or gibbon$1 OR owl$1 or avian or bird$1 or chick$3 OR goose or geese or duck$1 OR frog$1 or toad$1 or amphibian$1 OR fish$2 or zebrafish or marine or dolphin$1 OR in-vitro or invitro OR ecolog*) |
| PubMed | 01 January 2020 – 31 March 2021 | (((Axium) OR (AxiumTM) OR (AxiumDRG) OR (AxiumDRGTM) OR (AxiumTMDRG) OR (AxiumTMDRGTM)) AND ((Abbott) OR (Jude) OR (SJM) OR (dorsal root ganglion) OR (DRG) OR (neuromodul*) OR (neurostimulation) OR (neurostimul*) OR (neuro-modul*) OR (neuro-stim*))) AND (("2020/01/01"[Date - Publication] : "2021/03/31"[Date - Publication])) Filters: Humans, English |

## Post-Market Surveillance Data

The PMS data will be collected based on the PMS plan 90345660, Rev D. The PMS data generated from EU and world-wide complaints and reportable vigilance incidents associated with the Axium Neurostimulator System received through March 31, 2021 will be analyzed in the CER.

PMS data generated from corrective and preventative actions (CAPAs) and field safety corrective actions (FSCAs) associated with the Axium Neurostimulator System through March 31, 2021 will be analyzed in the CER.

## Post-Market Clinical Follow-Up Data

The PMCF studies has been conducted as per PMCF plan 90330520B. The Completed Post market studies involving Axium products are summarized in the current CER, CR002, Rev R and will also be included in the Rev S CER. Updates and any available data from the following active PMCF activities will be analyzed in the CER:

- TARGET (Targeted SCS (DRG) Stimulation) post-approval study (PAS; CRD 810; Protocol SJM-CIP-10113). The TARGET study is a post market study is designed to demonstrate continued safety of spinal cord stimulation of the DRG for the treatment of chronic, intractable lower extremity neuropathic pain associated with a diagnosis of CRPS and peripheral causalgia. NOTE: This study is a follow on from the ACCURATE study and includes data for both Axium and Proclaim devices.   Current AIMDD certified Axium devices Auxiliary Magnet Kit (MN23300) and Programmer Charger Kit (MN23400) are in the TARGET study.  The study includes US models of Axium devices.

- REALITY (CRD 959; Protocol ABT-CIP-10279 which includes the current AIMDD certified and commercially available Axium Clinician and Patient Programmers and other accessories.

# Appraisal of Pertinent Data for Clinical Evaluation (STAGE 2)

The relevant data sources identified in Stage 1 are appraised for suitability to provide clinical data relevant to the safety and performance of the Axium and weighted as described below. To assess contribution of evidence towards the clinical evaluation of safety and performance of the Axium, each clinical data set will be appraised in the CER in terms of scientific validity and relevance and weighted according to the pre-defined criteria described below.

### 8.1 Appraisal Criteria

Each of the identified clinical data sets will be appraised for their suitability and data contribution toward the clinical evaluation, according to the International Medical Device Regulators Forum (IMDRF) Medical Device Clinical Evaluation Working Group (MDCE WG) WG/N56FINAL:2019 guideline and MEDDEV 2.7/1 Rev 4. The appraisal criteria are presented in Table 19 .

Table 19. Appraisal Criteria for Suitability and Data Contribution

| Appraisal Criteria for Suitability |  |  |
| --- | --- | --- |
| Suitability Criteria | Description | Grading System |
| Appropriate Device | Were the data generated from the device in question? | D1 = Actual device D2 = Comparable device D3 = Other device |
| Appropriate Device Application | Was the device used for the same intended use (e.g., methods of deployment, application, etc.)? | A1 = Same Use A2 = Minor deviation A3 = Major deviation |
| Appropriate Patient Group | Were the data generated from a patient group that is representative of the intended treatment population (e.g., age, sex, etc.) and clinical condition (i.e., disease, including state and severity)? | P1 = Applicable P2 = Limited P3 = Different population |
| Acceptability of the Report / Data Collection | Do the reports or collations of data contain sufficient information to be able to undertake a rational and objective assessment? | R1 = High quality R2 = Minor deficiencies R3 = Insufficient information |
| Appraisal Criteria for Data Contribution |  |  |
| Data Contribution Criteria | Description | Grading System |
| Data Source Type | Was the design of the study appropriate? | T1 = Yes T2 = No |
| Outcome Measures | Do the outcome measures reported reflect the intended performance of the device? | O1 = Yes O2 = No |
| Appropriateness of Follow-up Period | Is the duration of follow-up long enough to assess whether duration of treatment effects and identify complications? | F1 = Yes F2 = No |
| Statistical Significance | Has a statistical analysis of the data been provided and is it appropriate? | S1 = Yes S2 = No |
| Clinical Significance of the Treatment Effect | Was the magnitude of the treatment effect observed clinically significant? | C1 = Yes C2 = No |
| Appraisal criteria defined per guidance in the IMDRF MDCE WG/N56FINAL:2019 Guidance, Appendix F, A Possible Method of Appraisal |  |  |

### 8.2 Analysis Criteria for Quality of Clinical Data

The appraisal criteria described in Section 8.1 will be used to identify data for analysis. As suggested by WG/N56FINAL:2019, the more Level 1 grades, the greater the weight of evidence provided in comparison to other data sets. As applicable: Equivalence of the devices has been established based on clinical, technical, and biological characteristics according to the MEDDEV 2.7/1 Rev 4 guidelines or EU MDR Annex XIV, Part A, Section 3.

The clinical data appraised for suitability and data contribution will be further assessed to evaluate their relative contribution towards safety and performance. The appraisal criteria described in Section above will be used to identify data for analysis. As suggested by WG/N56FINAL:2019, the more Level 1 grades, the greater the weight of evidence provided in comparison to other data sets.

In general, datasets will be weighted in the following manner:

Table 20. Level of Evidence Criteria

| Level of Evidence | Criteria |
| --- | --- |
| Level1 | Data on the <subject or equivalent device(s)>, [D1] used as intended [A1] in an applicable patient group [P1], from all study designs [T1 or T2], with high quality or minor deficiencies in acceptability / data collation [R1 or R2], outcome measures that reflect performance [O1], an appropriate follow-up period [F1], appropriate statistical analysis [S1], and a clinically significant effect [C1] were given the greatest weight in the evaluation [D1, A1, P1, T1/T2, R1 and R2, O1, F1, S1, C1]. |
| Level 2 | As it is important for manufacturers to understand off-label usage [A2, A3], use in patient groups with limited or no applicability [P2, P3], datasets pertaining to the subject devices [D1] identified during the search with high quality or minor deficiencies [R1 or R2], outcome measures that may reflect performance [O1], an appropriate follow-up period [F1], appropriate statistical analysis [S1], and a clinically significant effect [C1] were also be evaluated. To capture available data, these studies included all study design types [T1, T2, T3]. [D1/D2, A2/A3, P2/P3, T1/T2/T3, R1/R2, O1/O2, F1, S1, C1]. |
| Level 3 | Data from current knowledge/state of the art and similar competitor devices [D3] will also be analyzed, both on- and off-label [A1-A3] in all patient groups [P1-P3] with high quality or minor deficiencies [R1 or R2], all study designs [T1], outcome measures that reflect performance [O1], an appropriate follow-up period [F1], appropriate statistical analysis [S1], and a clinically significant effect [C1]. [D3, A1/A2/A3, P1/P2/-P3, T1, R1/R2, O1, F1, S1, C1]. |

Reports meeting these criteria will be critiqued and analyzed as part of the clinical evaluation.

# Analysis of Clinical Data (Benefit-Risk Analysis) (STAGE 3)

## Safety and Performance

The safety and performance measures defined in Section 3.0 for the Axium Neurostimulator system will be used in the CER to guide the analysis of clinical data to determine if the relevant ER are met.

Data addressing the safety and performance measures defined in Section 3.0 will be extracted from the identified sources of clinical data. The clinical data will be collectively analyzed to determine whether the Axium Neurostimulator system meet the pre-defined criteria for safety and performance.

## Risk Analysis (from Risk Management Documentation)

The risk management process is conducted according to procedure defined in 50-0043, Risk Management SOP (Standard Operating Procedure)’ and Risk Management Plan (RM004, Rev P) in accordance with principles of ISO14971:2012 Medical devices - Application of Risk Management to Medical Devices.

The purpose of the risk assessment is to identify known and foreseeable hazards, their potential adverse effects, and their corresponding potential causes. Risk is defined as the combination of the probability of occurrence of harm and the severity of that harm.

All Abbott products are reviewed for risk and undergo a Failure Modes and Effects Analysis (FMEA) and/or Hazard Analysis. Risk assessment reports are reviewed at defined time intervals and updated based on data from literature, commercial complaints, and clinical trials.  Risk reduction was considered for all hazards regardless of level; Low, Medium or High. The FMEA for the design, manufacturing processes, and use for the Axium neurostimulator System are presented in Table 21 after application of design, use error, and process mitigations. All residual risks identified as Low in the risk management reports are considered acceptable since the hazards have been mitigated as low as possible. The remaining Medium and High residual risks are associated with potential use errors, design, process, or material hazards.

No residual risks have been identified through Risk Management Report – DRG Leads, Extension and Accessories, 90248254 Rev. E, Cybersecurity Risk Assessment 90305433 Ver. B and Risk Management report, Spinal Modulation, INC, RM006 Rev H.Table 21. Residual Risks

| Device | Risk Documentation | Residual Risks |
| --- | --- | --- |
| DRG Leads, Extension and Accessories | 90248254 Rev. E | No residual risks |
| Spinal Modulation, INC | RM006 Rev H | No residual risks |
| Axium Neuromodulation System | 90305433 Ver. B | No residual risks |

## Benefit-Risk Assessment

All data that qualifies for evaluation per the appraisal methodology described in Section 8.0 will be evaluated as follows:

- Any new new/unknown AEs/patient harms/ device defects not in the current IFU or risk documentation will be identified and discussed in the context of clinical benefit-risk of the subject device

- Any safety or performance unfavourability reported for the subject device based on comparison between groups or cohorts will be identified and discussed in the context of clinical benefit-risk of the subject device

- Any significant or consistently observed deviations in rates of pre-determined safety and performance objectives will be discussed. The acceptance /threshold rates for these S&P (Safety and Performance) objectives will be used to make this determination. These rates will be based on SOTA rates and/or rates of OPCs defined in the ISO-Standard relevant for the subject device. These acceptance rates will be defined in the CER (or, if already defined remove this statement).

- Any residual risks identified through risk assessment activities will be discussed in the context of clinical benefit-risk (as applicable/ If applicable).

Based on the assessment above a conclusion will be made about the clinical benefit-risk of the subject device and its placement within the State-of-the-Art.

The Overall benefit-risk assessment statement will be presented in the CER.

## Additional Data

If the collective data from the above data sources are determined to be insufficient to make a reasonable conclusion regarding the acceptable safety and performance of the devices under consideration, the following actions/items will be performed:

- Identify any gaps or missing data

- Determine the additional clinical data and data source(s) that are needed so that conclusions can be drawn with confidence in relation to conformity with the essential requirements, including:

- Evaluation of the safety, performance and benefit/risk profile

- Compatibility with a high level of protection of health and safety

- The acceptability of any undesirable side-effects

- The risk of use error and the adequacy of the IFU to intended users

# Summary of the Clinical Evaluation (Stage 4)

The clinical evidence will be critically assessed and summarized in the CER to confirm whether the safety and performance of the medical devices under assessment, when used under the conditions and for the purposes intended, as specified by the manufacturer, are in compliance with Annex 1 of the AIMDD 90/385/EEC, MEDDEV 2.7.1, Revision 4 and ISO 14155:2011. Conclusions will be made in the CER with respect to the following:

- Compliance with ERs on clinical safety and performance of the devices under assessment when used according to their intended purpose, including the acceptability of the benefit/risk profile and of side-effects.

- Consistency and alignment between the clinical evaluation and the information materials supplied by the manufacturer (including the label, IFU, available promotional materials, including accompanying documents possibly foreseen by the manufacturer);

- Consistency between these informational materials and the current knowledge/ state of the art;

- Whether additional clinical investigations or other measures are necessary in order to generate additional data so that conclusions can be drawn regarding conformity with the ER;

- Residual risks and uncertainties or unanswered questions (including on rare complications, long term performance, safety under wide-spread use), whether these are acceptable for continued CE-mark certification, and whether they are required to be addressed during PMCF and /or PMS.

# Updating the Clinical Evaluation Plan

The CEP will be updated as needed throughout the lifecycle of the Axium.

The schedule of updates will be determined based on the conclusions in the CER and will be performed per the frequency specified in SOP 87302 at a minimum.

.

# APPENDICES

# Appendix 1: Axium models for US and Australia

The models in the Table 22 below are marketed and sold in the US and Australia.

- US Models: model numbers MN1XXXX

- Australia/New Zealand: model numbers ending with an “AU” suffix

- Slim Tip Leads: lead model numbers ending with an “A” suffix

- Ball Tip Leads: lead model numbers with no suffix (no longer manufactured)

Table 22.  Axium Models for US and Australia Markets

| Device Name | Description of Device | Model Numbers |  |
| --- | --- | --- | --- |
|  |  | United States | Australia |
| Trial Neurostimulator | Trial Pulse Generator | MN10100 | MN20100AU |
| Implantable Neurostimulator | Implantable Neurostimulator (IPG) including lead port plugs, torque wrench | MN10200 | MN20200AU |
| Trial SlimTip Lead Kit, 50cm and 90cm | Slim Tip Trial Lead including delivery sheaths, guidewire, stylets, delivery needle and tissue anchors | MN10350-50A, MN10350-90A | N/A |
| Implantable Lead (Ball Tip), 50cm and 90cm | Implantable Lead including delivery sheaths, guidewire, stylets, delivery needle and tissue anchors | N/A | MN20450-50AU, MN20450-90AU |
| Implantable SlimTip Lead, 50cm and 90cm | SlimTip Implantable Lead including delivery sheaths, guidewire, stylets, delivery needle and tissue anchors | MN10450-50A, MN10450-90A | MN20450-50AAU, MN20450-90AAU |
| Connector Cable Kit | Connector Cable, Lead Identifiers | MN11350 | MN21350AU |
| Tunneling Tool Kit, 30cm/51cm | Tunneling tool, straw, trocar and pencil tips, INS sizer, port plugs, torque wrench, hex key, sterile magnet sleeve | MN11900 | MN21900AU, MN22200AU |
| Clinical Programmer | Clinical Programmer, external magnet, charger, carrying case | MN10700 | MN20700AU |
| Patient Programmer | Patient Programmer, external magnet, charger, carrying case, medical alert card | MN10600-02 | MN20600AU, MN20600-02AU |
| Lead Accessory Kit | Delivery needles, tissue anchor, stylets, 30cm Big and Small Curve Delivery Sheaths | MN12050 | MN22050AU |
| 22cm Small Curve Delivery Sheath Kit | 22cm Small Curve Delivery Sheath | MN12150 | MN22150AU |
| 22cm Big Curve Delivery Sheath Kit | 22cm Big Curve Delivery Sheath | MN13650 | MN23650AU |
| 22cm Axium Small Curve Delivery Sheath Kit | 22cm Axium Small Curve Delivery Sheath | N/A | MN23850AU |
| 22cm Axium Big Curve Delivery Sheath Kit | 22cm Axium Big Curve Delivery Sheath | N/A | MN23950AU |
| Lead Extension Kit | Lead extension 50cm torque wrench | MN10550-50 | MN20550-50AU |
| Clinical Programmer Carrying Case | Programmer Carrying Case | MN13500 | N/A |
| Patient Programmer Carrying Case | Programmer Carrying Case | MN13500-S | N/A |
| Curved Needle | 4.5” 14G Curved delivery needle | MN14000 | MN24000AU |

# Appendix 2: Literature Search Protocol

# Background

The Clinical Literature Review, based on this protocol, will present data pertaining to the safety and performance of Abbott Medical’s Axium Neurostimulator System in the treatment of chronic intractable pain. The Axium Neurostimulator System components which are in scope of this evaluation are outlined in Table 1. The literature review will also provide data on current interventions (Includes therapies, diagnostic measures, measures for the management of diseases or medical conditions) for the intended patient population (state of the art) in order to give input to the assessments of acceptable benefit/risk profiles, what is currently considered as providing a high level of protection of health and safety and what are considered acceptable side-effects.

The Axium Neurostimulator System is similar to other competitive DRG systems and is comparable to other SCS neurostimulation systems (Nevro, Senza® and Medtronic, Recharging Restore™). Therefore, the literature review will identify and evaluate data from literature that studied both the safety of implanted neurostimulation systems regardless of indication and the safety of comparable implanted SCS systems. The previous literature search was conducted by manufacturer for the period of 01 December 2018 – 31 December 2019 and the results were summarized in CER CR002, Rev R.

# Objective of the Literature Search

The purpose of the systematic literature review is to provide a comprehensive analysis of the available published clinical literature to asses and analyze the body of evidence relating to clinical safety and performance data of Abbott Medical’s Axium Neurostimulator System.

The clinical literature review protocol and associated CER process is designed to meet the requirements as described in Annex 1: Essential Requirements), AIMDD 90/385/EEC, MEDDEV 2.7.1, Revision 4 and ISO 14155.

The key questions the literature review intends to answer are:

- What are the clinical outcomes associated with Abbott Medical’s Axium Neurostimulator System?

- What are the complications or adverse events associated with Abbott Medical’s Axium Neurostimulator System?

# Methods

A comprehensive literature search will be conducted by a qualified information analyst or technical librarian to identify published articles with pertinent clinical evidence that demonstrate the continued safety and performance of Abbott Medical’s Axium Neurostimulator System. The search will be for English language articles from peer-reviewed journals that provide data from clinical trials with human subjects. The clinical trials should evaluate the efficacy of Abbott Medical’s Axium Neurostimulator System and provide information regarding adverse events or complications during the study.

Table 23. Literature Search summary

| Date of Search | XXXX |
| --- | --- |
| Name of Person(s) Undertaking the Literature Search | XXXX |
| Period Covered by Search | January 01, 2020 – March 31, 2021. |
| Literature Sources Used to Identify Data and reason | The Allied & Complementary Medicine™, BIOSIS Previews®, Embase®, EMCare®, International Pharmaceutical Abstracts, MEDLINE®, Dimensions, PubMed Central and ToxFile® scientific literature bibliographic databases will be searched. Reason: These databases were selected to provide the most comprehensive coverage of the medical and biomedical literature and to ensure that all relevant studies will be identified while minimizing selection bias. |
| Additional search/ ad hoc searches/Internet search | In addition, other methods may be utilized such as “ad hoc searches” to adequately capture other relevant clinical literature. Articles that may have not appeared in the scientific literature bibliographic database results but provide pertinent information regarding Abbott Medical’s Axium Neurostimulator System safety or efficacy will be added. Articles may be identified ad hoc by reviewing the reference section of articles identified in the literature searches or through internet searches. |
| Additional search Tool | The results of the Allied & Complementary Medicine, BIOSIS Previews, Embase, EMCare, International Pharmaceutical Abstracts, MEDLINE, ToxFile, Dimensions and PubMed Central scientific literature bibliographic database searches will be imported into an internal literature management system (ProQuest Dialog), which allows searches on the full text, abstract, keywords, and title of those imported articles by the subject device and equivalent comparator brand names and manufacturers. ProQuest Dialog is an essential tool in identifying literature on the safety and performance of subject devices and equivalent comparators, because it allows extensive searching on all relevant articles based on its ability to pull full text search results into an environment where the full text can be searched for device names and provides full text articles (as available) from search results. In addition, reference lists from review articles will also be searched for additional relevant articles. |

Database Search Details

A comprehensive search strategy should include terms that cover five key areas. This formula ensures the identification of all relevant clinical safety and performance data on the subject device and/or equivalent comparator in the intended indication:

- Clinical conditions, disease, problem, indication

- Anatomical locations, intended population

- Interventions, therapies, generic device names

- Trade device names, manufacturers

- Other relevant terms

The search terms are detailed in Table 24. These preliminary search terms will be combined into technical database search terms and database queries.  Adverse events are not part of the search terms, because all clinical safety and performance outcomes will be sought in the target population irrespective of the measured or reported outcomes.

Table 24:  Preliminary Search Terms

| Search Term Categories | Specific Search Terms |
| --- | --- |
| Clinical Condition, Disease, Problem, Indication | Chronic Intractable Pain, Chronic Pain, Failed Back Surgery Syndrome, Intractable Low Back and Leg Pain, Intractable Low Back Pain, Intractable Leg Pain, Complex Regional Pain Syndrome Type I, Post Laminectomy Syndrome, Complex Regional Pain Syndrome Type II, Reflex Sympathetic Dystrophy, Intractable Neuropathic Pain, FBSS, CRPS |
| Anatomical Locations, Intended Population | Spine, Low Back, Leg, Limbs, Trunk, Dorsal Root Ganglia |
| Interventions, Therapies, Generic Device Names | Neuromodulation, Neurostimulation, Dorsal Root Ganglion Stimulation, DRG Stimulation |
| Trade Device Names, Manufacturers | SJM, St. Jude Medical, Abbott, Abbott Medical, Axium, Axium Neurostimulator System, Spinal Modulation Neurostimulator System |

Selection Criteria to Choose Articles

Titles and abstracts of references identified by the search strategy will be assessed in terms of their relevance and design according to the selection criteria outlined below in Table 25.

Table 25. Inclusion / Exclusion Criteria

| Exclusion Criteria | Inclusion Criteria |
| --- | --- |
| (E1) No abstract available (except clinical practice guidelines and expert consensus statements) (E2) Not published in English (E3) Duplicated references (E4) Preclinical studies (E5) Conference abstracts and (E6) Commentary, Review articles and meta-analyses, method papers, editorials, letters without original data and studies where clinical data are unextractable (E7) Article presents information covered in another included source (duplicate information) (E8) Article does not meet any inclusion criteria | (I1) Safety and Performance - Article reports on the use of the subject device in >1 patient - Article reports clinical outcomes on safety and performance of the subject device(s) when used as intended - Clinical outcomes reported are specific to the subject device(s), and results are not pooled with other devices (I2) PMS - Case reports or other articles describing the use of the subject device that don’t meet the criteria in (I1) (I3) State of the Art - Articles informing on the SOTA regarding DRG stimulation, the relevant patient population, or medical alternatives - Article reports human data relevant to the safety and performance of DRG stimulation (I4) Competitive / similar - Article reports benchmark clinical study data of a competitive device in >30 patients - Article reports clinical outcomes of safety and performance of a competitive device when used as intended Clinical outcomes reported are specific to the competitor device (i.e., results not pooled with other devices) |

# Data Collection Plan

Eligibility (Screening and Selection): The references identified by the search strategy will be assessed in terms of relevance and design according to the screening criteria above. All relevant publications will be reviewed and those considered eligible will be included.  Publications out of scope when assessed against the inclusion and exclusion criteria will be excluded with a reason for exclusion assigned.

Data Appraisal Plan

Each of the identified clinical data sets will be appraised for their suitability and data contribution toward the clinical evaluation of Axium Neurostimulator based on the criteria described in Section 8.1 The methodological quality of individual papers will be assessed based on the criteria described in Section 8.2.

Data Extraction

Data from the included studies will be extracted using predefined extraction tables. Once data extraction is completed, a table with all extracted data will be generated. The Data Extraction Table includes the following general elements for each included article:

- Reference

- Study design/evidence level

- Study population characteristics

- Clinical performance results

- Clinical safety outcomes

More specifically, the clinical performance and safety outcomes will include the variables or endpoints displayed in Table 26.

Table 26: Outcomes Collection Plan

| Variable | Variable Collected |
| --- | --- |
| Study Design | Study Objective(s) Study Design Number of Sites |
| Patient Characteristics | Diagnosis/Indications Device Type (Trial Lead, etc.) Device Brand Name Number of Patients Country Age Gender Follow-up Time |
| Performance Outcomes | 36-Item Short Form Health Survey (SF-36) EQ5D Pain Visual Analog Scale (VAS) Patient Reported Percent Pain Relief Patient Satisfaction Quality of Life Improvement McGill Pain Questionnaire Oswestry Disability Questionnaire Numeric Rating Scale Pain Catastrophizing Scale PROMIS-29 Analgesic Use Opioid Use |
| Safety Outcomes | Infection Undesirable Changes in Stimulation Pain, Redness or Numbness at the IPG Site Pain, Redness or Numbness at the Electrode Site Diminished or Loss of Pain Relief Lead Migration Battery Failure Cerebro Spinal Fluid Leakage Local Skin Erosion Seroma at Implant Site Wound Dehiscence Pulse Generator Migration Lead/Extension Breakage Reoperation Revision/Device Removal Death Other Adverse Events |

Performance outcomes and the numbers of patients will be extracted and recorded for the various study endpoint(s.) In general, outcomes will be extracted at the latest time point (i.e., longest duration of mean/median follow-up) to determine long term success. Multiple safety and performance variables will be considered for collection and analysis; however, it is not expected that all studies will have evaluated the same performance or safety outcomes.

All adverse events occurring in patients who received the subject device will be recorded. Device specific and procedure specific safety data will consist of all adverse events or complications that might or might not be related to the device, in the opinion of the investigator. When this information is not available, the data will be extracted as reported by the authors, including the numbers of patients evaluated for safety. The types and rates of safety events will be reported.

Data Analysis Plan

Clinical outcomes from individual, included articles will be assessed for device safety and performance as well as consistency with the State of the Art (taking into account variability in treatment populations and study protocols). The clinical data will not be pooled across articles; instead ranges and rates will be summarized in CER for individual publications.

# REFERENCES

- Deer TR, Levy RM, Kramer J, et al. Dorsal root ganglion stimulation yielded higher treatment success rate for complex regional pain syndrome and causalgia at 3 and 12 months: a randomized comparative trial. Pain. 2017;158(4):669-681.

- Deer TR, Pope JE, Lamer TJ, et al. The Neuromodulation Appropriateness Consensus Committee on Best Practices for Dorsal Root Ganglion Stimulation. Neuromodulation. 2019;22(1):1-35.

- Dworkin RH, Turk DC, Wyrwich KW, et al. Interpreting the clinical importance of treatment outcomes in chronic pain clinical trials: IMMPACT recommendations. J Pain. 2008;9(2):105-121.

- Hays RD, Spritzer KL, Fries JF, Krishnan E. Responsiveness and minimally important difference for the patient-reported outcomes measurement information system (PROMIS) 20-item physical functioning short form in a prospective observational study of rheumatoid arthritis. Ann Rheum Dis. 2015;74(1):104-107.