Post-Market Surveillance Plan for

Spinal Modulation Neurostimulator System (Axium Neurostimulator System)

Approval:

| Role/Function | Printed Name | Signature and Date (ddMMMyyyy) |
| --- | --- | --- |
| Global Regulatory Operations-Clinical Evaluation (GRO-CE) /Author | Amy Liu | See electronic signature in Windchill |
| GRO-CE PM/Director | Sangeeta Mantripragada | See electronic signature in Windchill |
| Clinical Engineering/R&D | Tj Baggett | See electronic signature in Windchill |
| Medical Affairs | Allen Burton, MD | See electronic signature in Windchill |
| Regulatory Affairs | Jordan Hanson | See electronic signature in Windchill |
| Clinical Affairs | Edward Karst | See electronic signature in Windchill |

Table of Contents

1.0	Purpose	4

2.0	Scope	4

3.0	Reference Documents	5

3.1.	Applicable Documents	5

3.2.	Applicable Procedures	5

3.3.	External References	6

4.0	Device description	6

4.1.	Subject Device Description	7

4.1.1.	General Product Description 	7

4.1.2.	Intended Use	7

4.1.3.	Indications	7

4.1.4.	CE Mark Information and Date Device was Placed in EU Market	7

4.2.	Similar Products Available on the Market	7

5.0	METHODS AND PROCEDURES for Post-market surveillance	8

5.1.	Complaint and Vigilance Handling	9

5.1.1.	Trending	10

5.2.	Risk Management	10

5.3.	Corrective Actions/Field Safety Corrective Actions (FSCA)	10

5.4.	Feedback by Users, Distributors and Importers	11

6.0	POST-Market clinical follow-up	11

6.1.	PMCF Study	12

6.2.	New Clinical Study	12

6.3.	Device Registry / Hospital Registry	13

6.4.	Restrospective Data from Patients previously exposed to the Device	13

6.5.	Clinical Literature	13

6.6.	Customer Survey	13

7.0	ANALYSIS AND DOCUMENTATION OF the PMS	13

8.0	APPLICATION of THE PMS FINDINGS	14

Table of Tables

Table 1: Axium Products (Current AIMDD Certified)	4

Table 2: Axium Products (Rationalized and Sold in other Geographies)	5

Table 3: Regulatory Approval History (Current AIMDD Axium devices)	12

Table 4: Regulatory Approval History (Current AIMDD devices used with Proclaim DRG      System)	12

Table 5: Regulatory History of Rationalized Axium Devices	13

Table 6 Countries Where Axium Neurostimulation System Approved	13

Table 7 Severity, Occurrence, and Detectability Table	17

Table 8 Post-Market EU Studies	19

Table of Figures

Figure 1: Illustration of TNS and INS Components	8

Acronyms and Abbreviations

| Term | Definition |
| --- | --- |
| AIMDD | Active Implantable Medical Device |
| MDD | Medical Device Directive |
| MDR | Medical Device Regulation |
| CEP | Clinical Evaluation Plan |
| CER | Clinical Evaluation Report |
| PMS | Post market Surveillance |
| PMCF | Post Market Clinical Follow-up |
| MAUDE | Manufacturer and User Facility Device Experience |
| EU | European Union |

# Purpose

This document defines the plan for post-market surveillance (PMS) of Spinal Modulation Neurostimulator System / Axium Neurostimulator System (Axium).

In a manner proportionate to the risk class and appropriate for the type of device, a PMS Plan documents the activities to systematically and actively gather, record and analyze relevant data on quality, performance and safety of the device throughout its entire lifetime.  In so doing, draw the necessary conclusions to determine, implement and monitor any preventive and corrective actions in order to update its technical documentation and identify the need for any preventive and corrective action.

# Scope

This document covers all activities intended to gain post-market safety and performance data for Axium during its entire lifetime, through active and passive means in a manner that is proportionate to the risk class and appropriate for the device type.  This post-market surveillance plan also includes a PMCF plan.

Information for the collection and utilization shall include, in particular:

information concerning , serious incidents, including information from post-market surveillance reports or periodic safety update reports (PSURs), CAPAs and field safety corrective actions;

- records referring to non-serious incidents and data on any undesirable side-effects;

- information from trend reporting;

- relevant specialist or technical literature, databases and/or registers;

- information, including feedbacks and complaints, provided by users, distributors and importers; and

- publicly available information about similar medical devices.

This plan applies to the following products:

# Reference Documents

## Applicable Documents

| Number | Title |
| --- | --- |
| CR002 | Clinical Evaluation Report— Axium Neurostimulator System |
| RM004 | Risk Management Plan— Axium Neurostimulator System |
| RM006 | Risk Management Report— Axium Neurostimulator System |
| 90305281 | Risk Management Plan— Orion IPG with Artemis 3.7 |
| 68-4930 | Risk Management Report - Orion Neuro Hybrid IPG |
| 90207284 | Risk Management Plan-Orion DRG |
| 68-4931 | Risk Management Report - Orion MR Conditional Neurostimulation System |
| CL1003132 | SMI 01 & SMI 02 Combined Final Clinical Study Report |
| 90330519 | Clinical Evaluation Plan— Axium Neurostimulator System |
| 90458061 | Proclaim DRG PMS/PMCF Plan |
| 90488409 | Post-Market Surveillance Data Summary Report— Axium Neurostimulator System |

## Applicable Procedures

| Number | Title |
| --- | --- |
| 87692 | Post-Market Surveillance System |
| 87303 | PMS Plan Process |
| 87302 | Clinical Evaluation Process |
| 87325 | CN Corrective and Preventative Action (CAPA) |
| 87326 | CN Exception Management CAPA process Work Instruction |
| 87457 | CN Product Experience Handling |
| 88052 | Creating and Submitting a MDV Regulatory report Work Instruction |
| 87615 | Risk Management Divisional Policy |
| 88197 | Active, Literature Review Work Instruction |
| 86819 | Post-Market Feedback |
| 86068 | Field Action |
| 87443 | Complaint Trending Work Instruction |
| 88052 | Creating and Submitting a MDV Regulatory Report Work Instruction |
| 86697 | Packaged Device Labs |
| 85994 | Techical Evaluation Process |
| 50-0043 | SOP, Risk Management Process |
| 50-0103 | SOP, Complaint Trend Review |
| 50-0121 | WI, Application and Design FMECA |
| 50-0131 | SOP, Device Master Record (DMR) |
| 56-0228 | WI, Complaint Trend Review Data Definition |
| 56-0252 | Adverse Health Outcome (AHO) Work Instruction |
| 56-0266 | WI, Process FMECA |
| 56-0267 | Production / Post-Production Risk Management WI |
| 56-0268 | WI, Software FMECA |
| 56-0269 | WI, Hazard Analysis |
| 94-2232 | Severity, Occurrence and Detectability |

## External References

| Number | Title |
| --- | --- |
| MEDDEV 2.12-1 rev 8 | Guidelines on a Medical Device Vigilance System |
| MEDDEV 2.12-2 rev 2 | Guidelines on Medical Devices, Post Market Clinical Follow-up Studies, A Guide for Manufacturers and Notified Bodies |
| MEDDEV 2.7/1 rev 4 | Guidelines on Medical Devices, Clinical Evaluation: A Guide for Manufacturers and Notified Bodies Under Directives 93/42/EEC and 90/385/EEC |
| 90/385/EEC | Council Directive, Active Implantable Medical Devices Directive (AIMDD) |
| 93/42/EEC | Council Directive, Medical Device Directive (MDD) |
| EU MDR 2017/745 | Regulation (EU) 2017/745 of the European Parliament and of the Council on Medical Devices |
| ISO-13485:2016 | International Organization for Standardization 13485:2016 Medical Devices, Quality Management Systems, Requirements for Regulatory Purposes |
| ISO-14971:2012 | International Organization for Standardization 14971:2012 Medical Devices, Application of Risk Management to Medical Devices. |

# Device description

The device subject to this PMS Plan and similar products available on the market are described to allow a correct characterization of the performance of the device and a comparison to be made between the device and similar products available on the market.

## Subject Device Description

### General Product Description

The Axium Neurostimulator System is composed of the devices listed in Table 1, which comprise the two subsystems pictured in Figure 1. The Axium Neurostimulator System is a fully-implantable medical device that stimulates neural tissue to produce analgesia for the treatment of chronic pain.  The system is classified as an Active Implantable Medical Device (AIMD) per the Active Implantable Medical Device Directive (90/385/EEC).

Figure 1: Illustration of TNS and INS Components

(For Component Identification Only)

The following is a description of each of the components of the Axium Neurostimulator System:

- Trial Neurostimulator (TNS)

The external TNS device connects to the Trial Lead(s) or Lead Extensions and is worn by the patient for up to 30 days during the trial period. The TNS device has a belt clip for the patient’s convenience.

- Implantable Neurostimulator (INS)

The Axium Implantable Neurostimulator (INS) is a non-rechargeable, 4-channel electronic device. It uses microelectronic circuitry, powered by a hermetically sealed battery, to generate a pulsed waveform to stimulate neural tissue. The electronic circuitry and battery are housed in a hermetically sealed titanium case.

Each neurostimulator has a unique internal identifier that allows the physician to identify the type of device through an X-ray. The radiopaque identifier inside the case allows identification of both the device manufacturer and model number using standard x-ray equipment. For the Axium Neurostimulator, the code is SM001, which identifies Spinal Modulation as the manufacturer.

The INS is packaged in a sealed inner tray within a sealed outer tray.

- Implant Leads/Trial Leads/Lead Extension

The Lead Kits contain the Leads and the individual delivery devices that are required for their placement.

Implant / Trial Leads: The Leads are designed for percutaneous introduction into the body using a special needle and a set of custom delivery tools provided in their respective kits. Each Lead is fitted with four cylindrical electrodes spaced at equal intervals, which are intended to provide stimulation at the target dorsal root ganglion (DRG). The temporarily placed Trial Leads are intended for use for up to 30 days.

Lead Extension: The Lead Extension consists of a silicone port header that accepts the Axium Trial Lead and Implant Lead. It is intended to extend the length of the lead and provide a connection between the lead and the Connector Cable, or the lead and the Implantable Neurostimulator System. The Lead Extension is intended for chronic implantation as a component of the Axium Neurostimulator System.

- Connector Cable Kit

The Connector Cable connects the Leads or Lead Extension to the external TNS. The Connector Cable is packaged separately from the Lead and Lead Extension Kit. The Connector Cable includes a connector and 2 extension cables for use as needed.

The Lead Identifiers are small bent wires that can be inserted into the proximal stylet lumen in the leads and are individually colored to identify the leads.

- Lead Accessories

Small / Big Curve Delivery Sheaths are intended to allow passage of the lead percutaneously into the epidural space. The labeled length of the sheath is the distance from the hub to the pre-shaped tip and the length of the curve at the tip is approximately 2 mm for the Small Curve and approximately 8mm for the Big Curve.

Axium Small / Big Curve Delivery Sheaths are also intended to allow passage of leads percutaneously into the epidural space. Axium sheaths are internally reinforced with thin stainless steel braiding.

Complex Curve and Straight Stylets assist in steering and positioning the lead within the epidural space. These are included within the 50cm and 90cm SlimTip trial lead kits.

The 14G Delivery Needle is intended to access the epidural space, providing a conduit for lead, guidewire and delivery sheath placement. It is available as a straight needle or a curved needle. The Axium Curved Needle is available in a separate package and is uniquely identified by the colored hub. The 14G Delivery Needle is only available in the Implant Lead, Trial Lead, and Lead Accessories Kits.

The Guidewire is optional and is intended to verify that the needle is in the epidural space after using a loss of resistance technique.

Soft Tissue Anchors are intended to anchor the Lead in the soft tissue or on the skin surface proximal to the distal contacts of the Lead.  The soft tissues anchors are packaged with the leads, and included in the lead accessories kit.

- Implantation Tools

The tunneling tool is used to provide a conduit for the Trial Lead, Implant Lead, or Lead Extension to the INS or away from the midline of the spine. It is packaged with 2 exchangeable tips: a blunt pencil tip and a sharp trocar tip. A straw is slid over the tunneling tool and when the steel handle is removed, the straw provides the conduit for tunneling.

The INS Sizer is approximately the same size as the INS and allows the physician to properly size the INS pocket. The INS sizer is packaged with the Axium tunneling tool kits (MN21900, MN22200).

The port plugs are used to fill unused ports in the INS. The port plugs are packaged with the INS, but spare port plugs are also packaged with the Tunneling Tool Kits for the convenience of the physician.

- Additional Accessories

Sterile Magnet Sleeve: The magnet is placed in the sterile sleeve to allow it to be used during the implantation of the INS. The sterile magnet sleeve is packaged with the Axium tunneling tool kits (MN21900, MN22200).

Medical Alert Card identifies the patient as a user of the Neurostimulation System. The Medical Alert card is packaged with the INS and Patient Programmer.

Programmer Charger is to be used with the Clinical or Patient Programmers to charge the battery or allow use of the Programmers while plugged into standard electrical outlets.

Programmer Carrying Case protects the Programmers from liquid.

Auxiliary Magnet allows the user to turn the NS off or activates RF to allow the user to communicate with the NS.

Hex Key allows the user to release a set screw in the INS header or Lead Extension header that has been unscrewed too far. The Hex Key is not to be used to tighten the set screw against the lead. The hex key is packaged with the Axium tunneling tool kits (MN21900, MN22200).

Clinical Programmer and Patient Programmer

The Clinical Programmer is used to program the stimulation parameters for both the TNS and the INS. The instructions for programming the TNS and INS devices are the same. The Clinical Programmer is used by the physician or clinical staff.

The Patient Programmer allows the patient to adjust the stimulation settings of the TNS and INS devices within limits preset by the clinician. The Patient Programmer also allows the patient to turn stimulation off, if necessary.

### Intended Use

This Axium Neurostimulator System is designed to deliver low-intensity electrical impulses to nerve structures. The system is intended to be used with leads and associated extensions that are compatible with the system.

The Axium Neurostimulator System is intended to treat chronic intractable pain that has been unresponsive to other methods.  Such pain may be caused by neuropathic conditions such as Complex Regional Pain Syndrome (CRPS, formerly known as Reflex Sympathetic Dystrophy or causalgia), Failed Back Surgery Syndrome (FBSS) and peripheral neuropathy.

### Indications

The Axium Neurostimulator System is indicated for the management of chronic intractable pain.

### CE Mark Information and Date Device was Placed in EU Market

The Axium Neurostimulator System for the treatment of chronic intractable pain was approved for commercial distribution in Europe in November 2011. In addition, the Axium Neurostimulator System for the treatment of chronic, intractable pain of the trunk and/or limbs was approved for commercial distribution in Australia in 2013. In 2014, communications and manual updates were sent to the European and Australian regulators and physicians with regard to the appropriate technique for lead removal. The Axium Neurostimulator System was approved for commercial distribution in the United States in February 2016.  The device has not been withdrawn from marketing for any reason related to its safety or effectiveness.

The Axium Neurostimulator System was originally approved with Spinal Modulation (located in Menlo Park, California) as the legal manufacturer and manufacturing facility.  In May 2015, Spinal Modulation became a wholly-owned subsidiary of St. Jude Medical.  St. Jude Medical developed and received CE marking approval for the next-generation DRG system (Proclaim DRG) in September 2016.  The Proclaim DRG system utilizes some of the same DRG leads and accessories as the Axium Neurostimulator System.

In order to consolidate facilities, the DRG leads, accessories and Axium programmers were transferred from Spinal Modulation to St. Jude Medical’s facility in Plano, Texas as the legal manufacturer and manufacturing facility.  The Axium Implantable Neurostimulator, Trial Neurostimulator and other Axium-specific devices were not transferred to Plano as they are phasing out due to the Proclaim DRG system.  The Spinal Modulation facility ceased manufacturing and permanently closed its doors in December 2017.  All CE marked devices manufactured by Spinal Modulation were placed on the EU market for distribution and the Spinal Modulation CE certificate was cancelled in March 2018. Tables 4 and 5 outline the CE marking history and current certificates for the devices in the Axium Neurostimulator System.

Table 3: Regulatory Approval History (Current AIMDD Axium devices)

| Device Name | Model | GMDN | Date of First CE Mark | Certificate Number (SJM) | Certificate Number (Abbott) |
| --- | --- | --- | --- | --- | --- |
| Clinical Programmer | MN27000 | 61588 | 16 Nov 2011 | CE 639587 | CE 706798 |
| Patient Programmer | MN20600-02 | 61588 | 04 Mar 2014 |  |  |
| Auxiliary Magnet Kit | MN23300 | 57996 | 16 Nov 2011 |  |  |
| Clinical Programmer Charger Kit | MN23400 | 17115 | 16 Nov 2011 |  |  |
| Patient Programmer Charger | MN23400-U | 17115 | 04 Mar 2014 |  |  |
| Clinical Programmer Carrying Case | MN23500 | 12500 | 16 Nov 2011 |  |  |
| Patient Programmer Carrying Case | MN23500-S | 12500 | 04 Mar 2014 |  |  |

NOTE: Several DRG products are rebranded to Abbott Medical.  For the devices indicated, we have mirror CE certificates – one for Abbott Medical, one for St. Jude Medical to facilitate the rebranding transition.

Table 4: Regulatory Approval History (Current AIMDD devices used with Proclaim DRG System)

| Device Name | Model | GMDN | Date of First CE Mark | Certificate Number (SJM) | Certificate Number (Abbott) |
| --- | --- | --- | --- | --- | --- |
| Implantable SlimTip Lead, 50cm and 90cm* | MN20450-50A, MN20450-90A | 36007 | 11 Dec 2014 | CE 639589 | CE 691529 |
| Lead Extension Kit* | MN20550-50 | 36007 | 16 Nov 2011 |  |  |
| 22cm Small Curve Delivery Sheath Kit* | MN22150 | 44063 | 16 Nov 2011 |  |  |
| 22cm Axium Big Curve Delivery Sheath Kit* | MN23650 | 44063 | 16 Nov 2011 |  |  |
| Curved Needle* | MN24000 | 58293 | 04 Oct 2013 |  |  |
| Lead Accessory Kit* | MN22050 | 44063 | 16 Nov 2011 |  |  |
| Connector Cable Kit | MN21350 | 47143 | 16 Nov 2011 |  |  |

Table 5: Regulatory History of Rationalized Axium Devices

| Component |  | Model Numbers | GMDN Code | Date of First CE Mark | Current Certificate |
| --- | --- | --- | --- | --- | --- |
| Trial Neurostimulator |  | MN20100 | 36007 | 16 Nov 2011 | No, Certificate expired 05MAR2018 |
| Implantable Neurostimulator |  | MN20200 | 36007 | 16 Nov 2011 | No, Certificate expired 05MAR2018 |
| Trial SlimTip Lead Kit, 50cm and 90cm |  | MN20350-50A, MN20350-90A | 36007 | 11 Dec 2014 | No, Certificate expired 05MAR2018 |
| Implantable Lead (Ball Tip), 50cm and 90cm |  | MN20450-50, MN20450-90, MN20350-50, MN20350-90 | 36007 | 16 Nov 2011 | No, Certificate expired 05MAR2018 |
| Tunneling Tool Kit, 30cm/51cm |  | MN21900, MN22200 | 44063 | 16 November 2011 | No, Certificate expired 05MAR2018 |
| Patient Programmer |  | MN20600 | 61393 | 16 November 2011 | No, Certificate expired 05MAR2018 |
| 22cm Axium Small Curve Delivery Sheath Kit |  | MN23850 | 44063 | 04 September 2013 | No, Certificate expired 05MAR2018 |
| 22cm Axium Big Curve Delivery Sheath Kit |  | MN23950 | 44063 | 04 September 2013 | No, Certificate expired 05MAR2018 |

The Axium Neurostimulator System is currently approved to be marketed in the following countries in Table 6.  Please note that this table reflects the approval status for the Axium Neurostimulator System prior to the certificate cancellation in March 2018.  All CE marked inventory was placed on the EU market prior to cancellation of the certificate and the remaining inventory is being consumed.

Axium products marketed and sold only in the US and Australia are listed in Table 2.

Table 6 Countries Where Axium Neurostimulation System Approved

| EUROPE | GLOBAL |
| --- | --- |
| Austria, Belgium, Denmark, Finland, France, Germany, Ireland, Italy, Luxembourg, Netherlands, Norway, Spain, Sweden, Switzerland, United Kingdom | Australia, Hungary, Kuwait*, New Zealand, Qatar, South Africa, United States |

*Kuwait will accept Axium products shipped to Ministry of Health hospitals

# METHODS AND PROCEDURES for Post-market surveillance

This section describes methods and procedures for actively and systematically gathering, recording and analyzing relevant data on the quality, performance and safety of Axium  throughout its entire lifetime, and in so doing, draw the necessary conclusions to determine, implement and monitor any preventive and corrective actions in order to update its technical documentation and identify the need for any preventive and corrective action.

Based on existing or available information on the safety and performance of the device, specific activities appropriate for adequate post-market surveillance of the device safety and performance should be described.

- The post-market surveillance plan shall address the collection and utilization of available information, in particular:

  - information concerning serious incidents, including information from PSURs, and field safety corrective actions;

  - records referring to non-serious incidents and data on any undesirable side-effects;

  - information from trend reporting;

  - relevant specialist or technical literature, databases and/or registers;

  - information, including feedbacks and complaints, provided by users, distributors and importers; and

  - publicly available information about similar medical devices.

- The post-market surveillance plan shall cover at least:

- a proactive and systematic process to collect any information referred to in point (a). The process shall allow a correct characterization of the performance of the devices and shall also allow a comparison to be made between the device and similar products available on the market;

- effective and appropriate methods and processes to assess the collected data;

- suitable indicators and threshold values that shall be used in the continuous reassessment of the benefit- risk analysis and of the risk management as referred to in Section 3 of Annex I;

- effective and appropriate methods and tools to investigate complaints and analyze market-related experience collected in the field;

- methods and protocols to manage the events subject to the trend report as provided for in MDR Article 88, including the methods and protocols to be used to establish any statistically significant increase in the frequency or severity of incidents as well as the observation period;

- methods and protocols to communicate effectively with competent authorities, notified bodies, economic operators and users;

- reference to procedures to fulfil the manufacturers obligations laid down for PMS System, PMS Plan and PMS report/PSUR (can be fulfilled through Section 3.2, Section 7 and Section 8);

- systematic procedures to identify and initiate appropriate measures including corrective actions;

- effective tools to trace and identify devices for which corrective actions might be necessary; and

- a PMCF plan or a justification as to why a PMCF is not applicable (see Section 6)

(c) Activities pertaining to collecting each aspect itemized in (a) above may be described as subsections below (examples) with specifications itemized in (b) above.

## Complaint and Vigilance Handling

Product Experience handling procedures are performed according to SOP 87457. Monthly complaint trend assessment is performed in accordance with SOP 50-0103. The post market data risk acceptability is evaluated through risk management activities driven by SOP 50-0043, ‘Risk Management Process’, WI 56-0269, ‘Hazard Analysis’, WI 56-0121, ‘Design and Application FMECA’, WI 56-0268, Software FMECA, WI 56-0266, Process FMECA. The residual risk for all devices reviewed in the Post Market Surveillance Data section below is acceptable per WI, Production/Post-Production Risk Management, 56-0267.

Per Abbott risk management process (SOP 50-0043) complaints are categorized by severity and rate calculation is done for each severity category. For risk acceptability of the modes, the risk management process is established per Hazard Analysis Work Instruction (WI) 56-0269 which defines the risk acceptability for patient/user facing risks. As shown below, this takes into consideration the severity of the adverse events as well as the occurrence rate at which these events are reported. The acceptability limits are defined and documented in 94-2232 Rev. H -  Severity, Occurrence, and Detectability (see Table 7). Table 7 provides guidance on acceptable risk regions based on the severity and overall probability of occurrence of patient effect. For lower severities a higher overall probability of patient effect is acceptable.

- For lower severity such as Severity 0 (No adverse health consequences), no limit has been established.

- For Severity 1 (Temporary/reversible injury or impairment not requiring acute medical intervention) a higher overall probability of patient effect is acceptable. An occurrence rate of 10% or lower is deemed acceptable risk.

- For Severity 2 (Temporary/reversible injury or impairment requiring acute medical intervention), an occurrence rate of less than 1% is deemed acceptable risk.

- For Severity 3 (Permanent impairment of body function or permanent damage to a body structure), an occurrence rate of less than 0.1% is deemed acceptable risk.

- For Severity 4 {Life threatening (death has or could occur)}, an occurrence rate above the acceptable risk level documented in the respective risk documents will require an investigation and a resolution.

Note: Risks which are rated as Unacceptable are mitigated to an acceptable risk level by performing a Risk Benefits Analysis (RBA) and documenting the results in the risk documents.

Per 56-0252- Adverse Health Outcome (AHO) WI, the following define Severity 1 and 2 events.

- AHO Severity 1 events: Limited or minor events that are temporary/reversible injury or impairment that do not require acute medical intervention (self-limited.) [Note: typically related to loss of therapy which is recovered without surgical intervention.]

- AHO Severity 2 events: Moderate to Serious events that are temporary/reversible injury or impairment that requiring acute medical intervention (e.g. surgical intervention).

- AHO Severity 3 events: Severe or Critical events that involved permanent impairment of body function or permanent damage to a body structure, that is associated with a substantial disruption of a person’s ability to conduct normal life functions.

- AHO Severity 4 events: Life Threatening or Catastrophic events in which life is threatened and death has occurred or could occur.

Risk region acceptability in the Hazard Analysis WI were established based on review of clinical/ scientific literature and relevant post-market data to ensure that the regions reflect risks that are inherent to the use of neurostimulator systems.

Table 7 Severity, Occurrence, and Detectability Table

A need for action is based on the monthly complaint trend analysis (Complaint Trend Review SOP 50- 0103, section 8.0) and using established statistical analysis if the complaint rate is above the predetermined Upper Control Limit (UCL) as described in ‘WI Complaint Trend Review Data Definition’: 56-0228. Please refer to section 8.1 step 2 in 56-0228. Additionally, depending on the severity and frequency of occurrence Abbott may determine the need to act prior to the rate reaching 1%.

### Trending

Monthly trending of complaints, returns and product analysis results will be conducted per the Performance Review System.  The performance review system will generate a rate of occurrence for each complaint category (e.g., sensing issues, labeling issues, premature battery depletion) as well as findings in returned products (e.g., component hardware or software deficiencies).

Complaints or issues that are trending higher than expected, or that have unique situations or failure modes, may be escalated to a Performance Review Investigation (PRI) or directly to a Corrective and Preventive Action (CAPA) as determined by the Performance Review Team and/or CAPA Review Board. The decision whether to escalate a trending complaint is also based on product risk management, which provide expected probabilities of risks and guidance for escalation, respectively.

## Risk Management

Field event data concerning serious, non-serious, and undesirable side effects is summarized in the Post-Market Surveillance section of the Risk Management Report (RMR) XXXXX for <Product Name>.

Residual risk acceptability is documented in the Criteria for Risk Evaluation and Overall Benefit Risk Analysis sections of the RMR and/or SRA (insert document number(s) for the RMR and/or SRA of the products in scope), which conveys information on the following three conditions: 1.) Individual risk levels fall below defined threshold values, 2.) Risk have been reduced as far as possible, and 3.) The overall benefits offered by use of the device outweigh the overall risks.

Continuous reassessment of risk is conducted in accordance with triggering guidance provided in the Risk Management Plan and Report Work Instruction.

## Corrective Actions/Field Safety Corrective Actions (FSCA)

Any action taken to reduce a risk of death or serious deterioration in the state of health associated with the use of [product] shall be treated as Field Safety Corrective Action (FSCA). Users shall be informed of such actions via a Field Safety Notice and Competent Authorities shall be notified in compliance with MEDDEV 2.12/1.

Abbott implements processes for documenting and managing potential or identified nonconformities with materials, products, processes, equipment, and/or quality systems within the Corrective and Preventive Action (CAPA) System (87326), also known as the Exception Management Process.

Procedures for determination, planning and execution of a field action, including FSCA are described in SOP 86068.

## Feedback from Users, Distributors and/or Importers

Abbott has multiple processes established to monitor post-market performance of its products and detect quality issues.  Collection and assessment of post-market feedback (SOP 86819) from the field is an important part of these activities.  Currently, there are <no specific plans OR following plans> in addition the customer complaint handling, analysis and management as described above.

# POST-Market clinical follow-up

As a part of the Axium ’s PMS Plan, a PMCF is planned to proactively collect and evaluate clinical data from the use in or on humans of the CE-marked Axium.

The PMCF Plan will document methods and procedures for collecting and evaluating clinical data with the aim of confirming the safety and performance throughout the expected lifetime of the device, of ensuring the continued acceptability of identified risks and of detecting emerging risks on the basis of factual evidence.

All Abbott PMCF plans specify the methods and procedures for proactively collecting and evaluating clinical data with the aim of:

- Confirming the safety and performance of the device throughout its expected lifetime;

- Identifying previously unknown side-effects and monitoring the identified side-effects and contraindications;

- Identifying and analyzing emergent risks on the basis of factual evidence;

- Ensuring the continued acceptability of the benefit-risk ratio; and

- Identifying possible systematic misuse or off-label use of the device, with a view of verifying that the intended purpose is correct.

## PMCF Study

In the EU, seven studies were carried out concurrently to ACCURATE and are summarized in Table 1.

Table 1 Post-Market EU Studies

| Study | Summary | Milestones | Estimated Timeline | References |
| --- | --- | --- | --- | --- |
| 04-SMI-2012 | 49 subjects with chronic intractable pain were followed for 12 months. VAS decreased from 8.1 to 4.1. One SADE was reported. | Enrollment began Apr 2012 | Last subject exited Mar 2015 | CL1002544 |
| 08-SMI-2012 PREDICT | 74 subjects with chronic intractable pain were followed to 24 months. VAS decreased from 7.5 to 4.4. Four SADEs were reported. | Enrollment began Nov 2012 | Last subject exited Jan 2017 | CL1003861 |
| 10-SMI-2012 CRD_854 | 16 subjects with Peripheral Neuropathy were followed to 12 months. VAS decreased from 7.8 to 3.5. Two SADEs were reported. | Enrollment began Jul 2013 | Last subject exited Oct 2016 | CL1003301 |
| 11-SMI-2013 BINGO | 18 subjects with chronic intractable pain were followed up to 12 months. VAS decreased from 7.6 to 4.8. 1 procedure-related SAE was reported. | Enrollment began May 2013 | Last subject exited Oct 2015. | CL1002982 |
| 18-SMI-2013 CRD_838 SYMPATHY | Nine subjects with chronic pain following surgical lumbar discectomy were followed to 12 months. VAS decreased from 8.4 to 2.9. One SAE not device related was reported. | Enrollment began Mar 2014 | Last subject exited Sep 2016 | CL1003222 |
| 20-SMI-2013 CRD_852 | 22 subjects with chronic post-surgical pain to were followed for 12 months. 24-month data was collected were appropriate. Four SAEs were reported, 1 of which was device related. VAS decreased from 7.6 to 4.9. | Enrollment began May 2013 | Last subject exited Dec 2016 | CL1003359 |
| 28-SMI-2015 CRD_849 | 14 subjects with discogenic low back pain were followed for 12 months. VAS decreased from 7.2 to 2.3. No SADEs were reported. | Enrollment began Oct 2015 | Last subject exited Jan 2018 | CL1004780 |

Study 04-SMI-2012

This study was a prospective, single-arm, multi-center post market study designed to assess the clinical effects of the commercially available Axium Neurostimulator System in the management of chronic intractable pain. Subjects were enrolled by two EU sites (n=66), 56 received a permanent implant and 49 were followed up to 12 months. The main etiologies represented FBSS, causalgia and CRPS-I.

The primary efficacy analysis measured the average pain reduction experienced using a Visual Analogue Scale (VAS) at 3, 6 and 12 months post permanent implant. Overall pain decreased from 7.4 cm at baseline to 4.7 cm at 12-months. Average VAS scores for the primary region of pain reported decreased from 8.1 cm to 4.1 cm.  In addition, an evaluation of the quality of life of subjects as measured by the EQ-5D assessment and an evaluation of the physical functioning of subjects using the Brief Pain Inventory assessment (BPI) was made comparing baseline data to 12 months post treatment data. Quality of life increased from 0.4 to 0.6. Pain severity decreased from 6.7 to 5.0 and interference decreased from 4.9 to 3.6. Adverse event data was collected over the entire period of the study. One SADE was reported.

Study 08-SMI-2012 (PREDICT)

The objective for this prospective, multi center DRG stimulation trial was to assess the performance of the commercially available CE marked Axium® Neurostimulator System in the management of chronic intractable pain. Five sites in the UK enrolled 135 subjects and 112 received a permanent implant. Twelve-month follow-up is available for 89 subjects and 74 are follow to 24-months.

The primary objective was pain relief measured using a Visual Analogue Scale (VAS). The average pain reduction experienced as well as the percentage of subjects with at least a 50% pain reduction were determined. VAS decreased from 7.5 cm to 4.4 cm at 24-months. 40% experienced at least a 50% in pain score. Secondary objectives included the occurrence of SAEs (n=4) and AE and quality of life measurements. Pain severity decreased from 6.9 to 4.8 at 24-months. Interference reduced from 7.1 to 5.0.

Study 10-SMI-2012 (DE Peripheral Neuropathy Study)

This was a prospective, single-arm, single center, post market study to assess the clinical effects of the commercially available Spinal Modulation Neurostimulator System in the management of peripheral neuropathy. A German site enrolled 33 subjects and 27 received a permanent implant. Sixteen subjects were followed for 12 months.

The primary objective of the study was pain relief as measured by the Visual Analog Scale (VAS). The average pain reduction and the percentage of subjects with at least a 50% pain reduction were also determined. Average pain decreased from 7.4 at baseline to 3.8 at 12-months. Pain severity decreased from 6.6 to 3.4 and interference from 5.4 to 3.6. SAEs were monitored and seven were reported.

Study 11-SMI-2013 (BINGO)

The objective of this Nordic multi-center, post market study was to assess the performance of the commercially available Axium Neurostimulator System in the management of chronic intractable pain. Two sites enrolled 41 subjects, implanted 31 with permanent devices and followed 18 subjects to 12 months. Causalgia was the main etiology enrolled. The performance was measured using pain relief (VAS score). Mean baseline VAS reduced from 7.6 to 4.8. Quality of life (EQ-5D) showed improvement across all dimensions measured. Pain severity decreased from 6.6 to 5.0 and interference decreased from 6.0 to 4.5. Safety was evaluated by monitoring the occurrence of SAEs and AEs throughout the study. One SADE was reported.

Study 18-SMI-2013 (SYMPATHY)

The objective of this multi-center, post market study was to assess the performance of the commercially available Axium Neurostimulator System in the management of chronic pain after surgical lumbar discectomy. Four sites enrolled 22 subjects, implanted 13 with permanent devices and followed 9 subjects to 12 months. The performance was measured using pain relief (VAS score). Mean baseline VAS reduced from 7.7 to 2.4. Quality of life (EQ-5D) showed improvement across all dimensions measured. Pain severity decreased from 7.3 to 2.8 and interference decreased from 5.7.0 to 2.3. Safety was evaluated by monitoring the occurrence of SAEs and AEs throughout the study. No SADEs were reported.

Study 20-SMI-2013 (NL Post-Surgical Study)

This post market observational study was a prospective, single-arm, multi-center study designed to assess the clinical effects of the commercially available Spinal Modulation DRG Neurostimulator System, in the management of chronic post-surgical pain. Subjects (n=30) were enrolled in two European sites and 26 received a permanent implant. Twenty-three subjects were followed up to 24 months. The primary objective was to measure pain relief using the VAS scale at baseline and follow-ups. Average overall pain decreased from 7.6 to 4.1 at 12 months and 4.9 at 24 months.  Pain severity on BPI decreased from 6.9 to 4.6 and interference from 4. 4 to 3.7. One SADE was reported.

Study 28-SMI-2015 CRD_849 (Discogenic Pain Study)

This post market observational study was a prospective, single-arm, observational study designed to investigate the effect of dorsal root ganglion (DRG) stimulation for a well-selected group of patients with discogenic lower back pain (LBP) with no history of previous back surgeries. Twenty subjects with confirmed discogenic LBP and no prior history of back surgery underwent trials of DRG stimulation and, if successful with at least 50% pain reduction, were permanently implanted. Subjects rated their pain, disability, quality of life, and mood at baseline, and 14 subjects were followed through 12 months of treatment. Study results showed DRG stimulation treatment for discogenic LBP improved the level of pain, function, and quality of life. Treatment with DRG stimulation reduced LBP ratings (68.3% reduction) after 12 months, and Oswestry ratings of disability significantly decreased after six months of treatment. During the temporary trial, three adverse events (AEs) were reported in three subjects: lead migration, changes in sensation related to stimulation, and disconnection of the external trial stimulator. All issues were resolved during routine placement of the permanent systems. After permanent implantation, ten AEs were reported in nine subjects. Four subjects had lead migration of which three resolved after surgical revision; a fourth subject was, by patient’s choice, explanted and withdrawn from the study. Three subjects reported a temporary return of their original pain level. This was found to be due to increased lead impedance and this was resolved through lead revisions (two subjects) or turning off the affected lead (one subject). Three subjects experienced pain at the subcutaneous neurostimulator pocket. In all cases, the event resolved after surgical revision. There were no serious AEs.

Table 7. PMCF Studies

| Study | Size | Clinical Endpoints | Duration/Status |
| --- | --- | --- | --- |
| 04-SMI-2012 | 66 | Average pain reduction and patients with ≥50% pain reduction using VAS | Completed |
| 10-SMI-2013 | 33 | Pain relief; Average pain reduction and patients with ≥50% pain reduction using VAS | Completed |
| 11-SMI-2013 (BINGO) | 41 | Pain relief (VAS); QOL (EQ-5D); physical functioning (BPI); pain distribution | Completed |
| 20-SMI-2013 | 30 | Pain relief (VAS) | Completed |
| TARGET Postapproval Study (PAS) | 376 | SAEs rate at 12 months; Change in overall pain intensity at 12 mos using VAS; Change in physical function using PROMIS scale; Change in quality of life using PROMIS scale | Follow-up: 1,3,6 and 12 months post-permanent implant |
| 08 UK Multi-Centre | 135 | Pain relief using VAS; adverse events; quality of life | Completed. Final Report in progress |
| 12 NL Multi-Centre | 5 | Change in perceived stimulation intensity during posture changes; devie activations per day; % of painful area covered by paresthesia; % of non-painful area covered by paresthesia; adverse events | Completed. Terminated early due to slow enrollment |
| 28 NL Multi-Centre | Data pending | Pain relief (VAS) | Data pending |

Table 8: PMCF Reports

| Study | Size | Clinical Endpoints | Report Number |
| --- | --- | --- | --- |
| 04-SMI-2012 | 66 | Average pain reduction and patients with ≥50% pain reduction using VAS | CL1002544 (Complete) |
| 10-SMI-2013 | 33 | Pain relief; Average pain reduction and patients with ≥50% pain reduction using VAS | CL1003301 (Complete) |
| 11-SMI-2013 (BINGO) | 41 | Pain relief (VAS); QOL (EQ-5D); physical functioning (BPI); pain distribution | CL1002982 (Complete) |
| 20-SMI-2013 | 30 | Pain relief (VAS) | CL1003359 (Complete) |
| TARGET Postapproval Study (PAS) | 376 | SAEs rate at 12 months; Change in overall pain intensity at 12 mos using VAS; Change in physical function using PROMIS scale; Change in quality of life using PROMIS scale | CL1003672 (current interim report) |
| 08 UK Multi-Centre | 135 | Pain relief using VAS; adverse events; quality of life | CL1003861 (Draft) |
| 12 NL Multi-Centre | 5 | Change in perceived stimulation intensity during posture changes; devie activations per day; % of painful area covered by paresthesia; % of non-painful area covered by paresthesia; adverse events | CL1003617 (Complete) |

## New Clinical Study

### REALITY

REALITY (CRD 959; Protocol ABT-CIP-10279): Long-Term Real-World Outcomes Study on Patients Implanted with a Neurostimulator is a new international clinical study.  The REALITY study is a post-market, prospective, non-randomized, single-arm, open-label study. This study is designed to collect active post-market surveillance data on Abbott’s market-released neurostimulation systems (including the MR-Conditional Proclaim DRG system) in accordance with the EU-MDR directive (2017/745). It is intended to collect short- and long-term safety and effectiveness data on various populations implanted with Abbott’s neurostimulation systems. This study has broad inclusion criteria and minimal exclusion criteria to ensure the results are representative of the real-world use of these devices. Enrollment caps will be implemented to ensure patients from approved indications are represented.

This study also includes the current AIMDD certified and commercially available Axium Clinician and Patient Programmers and other accessories. See Appendix section for all Axium models included in the REALITY PMCF study.  The study also includes the Axium models currently under used with Proclaim DRG. These are denoted with an * in the Appendix section Table 9.

The study will enroll up to 2,000 subjects from up to 100 participating centers in the U.S, Canada, Europe, Latin America, Australia, and emerging markets. Individuals who are scheduled to receive an implantable Abbott neurostimulation system (including MR-Conditional Proclaim DRG) are eligible for study consideration. Subject enrollment is expected to be completed within 4 years; subjects will be followed for 5 years. he total duration of the study is expected to be 10 years, including enrollment, data collection from all subjects, and study close out.  This study began enrolling in May 2019.

Clinical improvement from baseline to each follow up visit assessed by:

- PROMIS-29

- Numeric Rating Scale (NRS) for pain intensity

- Pain Catastrophizing scale (PCS)

- Oswestry Disability Index (ODI)

- Pain condition-related medication usage

- Transcutaneous Oxygen pressure measurement (TcPO2)

### TARGET STUDY

The TARGET (Targeted SCS (DRG) Stimulation) post-approval study (PAS; CRD 810; Protocol SJM-CIP-10113) is a prospective, non-randomized, single-arm study designed to demonstrate continued safety of spinal cord stimulation of the dorsal root ganglion (DRG) for the treatment of chronic, intractable lower extremity neuropathic pain associated with a diagnosis of complex regional pain syndrome (CRPS) and peripheral causalgia. NOTE: It is a follow on from the ACCURATE study and includes data for both Axium and Proclaim devices.

Current AIMDD certified Axium devices Auxiliary Magnet Kit (MN23300) and Programmer Charger Kit (MN23400) are in the TARGET study.  The TARGET follow-up details are documented in the Proclaim DRG PMS Plan (90458061).

The primary objective of this study is to demonstrate that the proportion of serious adverse effects (SAEs) at 12 months for subjects who receive the permanent Implantable Pulse Generator (IPG) is lower than a pre-specified objective performance goal. Secondary objectives will evaluate the change in overall pain intensity, change in physical function and change in quality of life.

The clinical study will be conducted at up to 45 centers in the United States. A maximum of 426 subjects will be enrolled to achieve 319 subjects with a permanently implanted system, which will account for expected attrition, to yield 287 subjects for the primary analysis at 12 months post-implant. Follow-up visits will occur at 1, 3, 6 and 12 months post-permanent implant. Additional data will be collected at 18- and 24-month visits, as applicable, for subjects who consent to the additional follow-up. This study began enrolling in August 2016. Enrollment is expected to be complete in December 2019. Twelve-month follow up data is expected in February 2021.

The following information will be collected at these visits:

- Visual Analog Scale (VAS) score

- Numeric Rating Scales (NRS)

- PROMIS-29 Profile domain scores

- PROMIS Global Health Scale score

- PGIC score

- Change in occupational status

- Change in activity level

- Satisfaction with device/therapy

- Willingness to undergo the procedure again

- Willingness to recommend the procedure to someone

- Sensations generated by stimulation

- Programming parameters

- AEs (if applicable)

- Protocol Deviations (if applicable)

## Clinical Literature

Review of published literature, including independently sponsored clinical studies, will occur via periodic Clinical Evaluation Report (CER) updates. The periodic CER updates are governed by the Clinical Evaluation Process SOP. Clinical literature is also reviewed per SOP 88197.

## Customer Survey

There are no active or planned customer surveys for the products in scope of this PMS Plan.

# ANALYSIS AND DOCUMENTATION OF the PMS

As required by risk management, post-market surveillance results may trigger an additional risk/benefit analysis, post-market investigations, and/or post-market corrections per the Corrective and Preventive Action (CAPA) procedures.

Each post-market activity outlined in this plan drives individual results reporting and monitoring (e.g., periodic trend reviews, risk management reviews). This includes a post-market clinical follow-up evaluation report (87312). The totality of all results of the post-market activities outlined in this plan are evaluated and documented annually in a PSUR.

# APPLICATION of THE PMS FINDINGS

The data from the post-market surveillance system will be used:

- To update the benefit-risk determination and to improve the risk management described in Risk Management SOP 50-0043 and the Risk Management Plans for products as described in RM004 and the Proclaim DRG System associated Risk Management Plans (90305281 and 90207284) respectively

- To update the design and manufacturing information, the instructions for use and the labelling as described in the Device Master Record SOP-50-0131, Packaged Device Labels-86697 and Technical Publications IFU-86944;

- To update the clinical evaluation per SOP 87302 – Clinical Evaluation Process;

- For the identification of needs for preventive, corrective or field safety corrective action per Corrective and Preventative Action (CAPA) SOP - 87325;

- For the identification of options to improve the usability, performance and safety of the device per Product Experience Handling SOP 87457;

- To contribute to the post-market surveillance of other devices when relevant per SOP 87457 – Product Experience Handling; and

- To detect and report trends per Complaint Trending Work Instruction (87443) and SOP, Complaint Trend Review (50-0103)