Post-Market Clinical Follow-up Plan

Axium Neurostimulator System

Approval:

| Role/Function | Printed Name | Signature and Date |
| --- | --- | --- |
| Global Regulatory Operations-Clinical Evaluation (GRO-CE) |  |  |
| Clinical Affairs |  |  |
| Clinical Engineering |  |  |
| Regulatory Affairs |  |  |

# Documentation Information

| Post-Market Clinical Follow-up (PMCF) Plan |  |  |  |  |
| --- | --- | --- | --- | --- |
| PMCF Plan Number: |  |  |  |  |
| PMCF Plan Date: |  |  |  |  |
| PMCF Plan Version: |  |  |  |  |
| Revision History |  |  |  |  |
| Rev. | Revision Date |  | Description of Change | Revised By |

# Table of Contents

A.	Manufacturer Contact Details	7

B.	Device Description and Specification	8

C.	Activities Related to PMCF:  General and Specific Methods and Procedures	10

C.1	Post-Market Study	11

C.1.1.	Basis/Need	12

C.1.2.	Description of the Activity	12

C.1.3.	Aim	13

C.1.4.	Rationale	14

C.1.5.	Known Limitations	14

C.1.6.	Timeline	14

C.2	Registry	14

C.2.1.	Basis/Need	15

C.2.2.	Description of the Activity	15

C.2.3.	Aim	15

C.2.4.	Rationale	15

C.2.5.	Known Limitations	15

C.2.6.	Timeline	16

C.3	Real World Data Analysis	16

C.3.1.	Basis/Need	16

C.3.2.	Description of the Activity	16

C.3.3.	Aim	16

C.3.4.	Rationale	16

C.3.5.	Known Limitations	16

C.3.6.	Timeline	17

C.4	Survey from Health Care Professional	17

C.4.1.	Basis/Need	17

C.4.2.	Description of the Activity	17

C.4.3.	Aim	17

C.4.4.	Rationale	17

C.4.5.	Known Limitations	17

C.4.6.	Timeline	17

C.5	Survey from Patients / Users	18

C.5.1.	Basis/Need	18

C.5.2.	Description of the Activity	18

C.5.3.	Aim	18

C.5.4.	Rationale	18

C.5.5.	Known Limitations	18

C.5.6.	Timeline	18

C.6	Case Reports Review	19

C.6.1.	Basis/Need	19

C.6.2.	Description of the Activity	19

C.6.3.	Aim	19

C.6.4.	Rationale	19

C.6.5.	Known Limitations	19

C.6.6.	Timeline	19

C.7	Scientific Literature Screening	20

C.7.1.	Basis/Need	20

C.7.2.	Description of the Activity	20

C.7.3.	Aim	20

C.7.4.	Rationale	20

C.7.5.	Known Limitations	20

C.7.6.	Timeline	20

C.8	Other Sources of Clinical Data	21

C.8.1.	Basis/Need	21

C.8.2.	Description of the Activity	21

C.8.3.	Aim	21

C.8.4.	Rationale	21

C.8.5.	Known Limitations	21

C.8.6.	Timeline	21

D.	Evaluation of Clinical Data Relating Equivalent or Similar Devices	22

E.	Reference to the Relevant Parts of the Technical Documentation	23

F.	Reference to Any Applicable Common Specifications, Harmonized Standards or Guidance Documents	24

G.	Estimated Date of the PMCF Evaluation Report	25

# Tables

Table 1: Description of Devices	23

# Figures

Figure 1: Subject flow through TARGET study.	31

Figure 2: REALITY Study Plan.	50

# Manufacturer Contact Details

| Legal manufacturer name | Abbott Medical |
| --- | --- |
| Address |  |
| SRN |  |
| Person responsible for regulatory compliance |  |
| E-mail |  |
| Phone |  |
| Fax |  |
| Authorized representative (if applicable) |  |
| Contact person |  |
| E-mail |  |
| Phone |  |
| Fax |  |

# Device Description and Specification

| Product or trade name | Axium Neurostimulator System |
| --- | --- |
| Model and type | DRG Neurostimulation System |
| General description of the device | The DRG Neurostimulator System is a fully implantable medical device that stimulates neural tissue to produce analgesia for the treatment of chronic pain. The system consists of an IPG and leads, accessories, programmers, and extensions that are compatible with the system. |
| Intended purpose | This Axium Neurostimulator System is designed to deliver low-intensity electrical impulses to nerve structures. The Axium Neurostimulator System is intended to treat chronic intractable pain that has been unresponsive to other methods |
| Intended users | Implanting physicians should be experienced in the diagnosis and treatment of chronic pain syndromes and have undergone surgical and device implantation training. |
| Basic UDI-DI | N/A |
| Intended patient population | The intended target population for the system are individuals with chronic, intractable pain or complex regional pain syndrome (Type 1 and Type 2), leading to moderate to severe chronic intractable pain of the lower limbs. |
| Medical condition(s) | Chronic intractable pain |
| Indications | The Axium Neurostimulator System is indicated for the management of chronic intractable pain. |
| Contraindications | Patients contraindicated for the Axium Neurostimulator System are those who: - Have an active implantable medical device including but not limited to cardiac pacemakers and cardiac defibrillators. - Are unable to operate the system. - Are poor surgical risks. - Are pregnant. - Are under the age of 18. |
| Warnings |  |
| List and description of any variants and / or configurations covered by this plan |  |
| List of any accessories covered by this plan | - Lead Accessories Kit (MN12050) - Tunneling Tool Kit (MN11900) - Tunneling Tool (Model MN15000) |
| Certificate number |  |
| CND code(s) |  |
| Class |  |
| Classification rule |  |
| Expected lifetime |  |
| Novel product | ☐ yes ☐ no |
| Novel related clinical procedure | ☐ yes ☐ no |
| Explanation of any novel features |  |

# Activities Related to PMCF:  General and Specific Methods and Procedures

A summary table of the different PMCF activities foreseen by the manufacturer is provided below:

| Number | Description | Aim | Rationale and known limitations | Timelines |
| --- | --- | --- | --- | --- |
| C1 |  | The primary objective of this study is to demonstrate that the proportion of serious adverse effects at 12 months for subjects who receive the permanent Implantable Pulse Generator (IPG) is lower than a pre-specified performance goal. |  | 70 Month |

Post-Market Study -TARGET

| Title |  | TARGET (DRG) Post-Approval Study |
| --- | --- | --- |
| Purpose |  | The ACCURATE study evaluated spinal column stimulation of the dorsal root ganglion (DRG) with the Axium Neurostimulator System for the treatment of chronic, intractable lower extremity neuropathic pain associated with a diagnosis of complex regional pain syndrome (CRPS) and peripheral causalgia (PC). The ACCURATE study met its primary endpoint and the results of the study provide reasonable assurance of safety and effectiveness of the Axium Neurostimulator System. The Axium Neurostimulator System was approved for sale in the United States by FDA and is indicated for spinal column stimulation via epidural and intra-spinal lead access to the dorsal root ganglion as an aid in the management of moderate to severe chronic intractable* pain of the lower limbs in adult patients with Complex Regional Pain Syndrome (CRPS) types I and II.** The Proclaim Neurostimulation System was subsequently approved for the same indication and is replacing the Axium device in commercial distribution. This post-approval study is a condition of FDA approval. *Study subjects from the ACCURATE clinical study had failed to achieve adequate pain relief from at least 2 prior pharmacologic treatments from at least 2 different drug classes and continued their pharmacologic therapy during the clinical study. **Please note that in 1994, a consensus group of pain medicine experts gathered by the International Association for the Study of Pain (IASP) reviewed diagnostic criteria and agreed to rename reflex sympathetic dystrophy (RSD) and causalgia, as complex regional pain syndrome (CRPS) types I and II, respectively. |
| Primary Objective |  | The primary objective of this study is to demonstrate that the proportion of serious adverse effects (SAEs) at 12 months for subjects who receive the permanent Implantable Pulse Generator (IPG) is lower than a pre-specified objective performance goal. |
| Secondary Objective |  | · To evaluate change in overall pain intensity. · To evaluate change in physical function . · To evaluate change in quality of life |
| Primary Endpoint |  | The primary endpoint is the 12-month SAE rate for subjects receiving a permanent IPG |
| Secondary Endpoint |  | Percent change from baseline to 12 months post-permanent implant for overall pain intensity measured using the Visual Analog Scale (VAS) · Change from baseline to 12 months post-permanent implant for physical function measured using the PROMIS-29 Profile. · Change from baseline to 12 months post-permanent implant for quality of life measured using the PROMIS Global Health Scale |
| Descriptive Endpoints |  | Safety · Proportion of subjects with device, procedure or stimulation related SAEs and non-SAEs among subjects who receive the permanent implant · Proportion of subjects who have AEs related to the trial implant procedure among subjects who fail the trial neurostimulator phase · Summary of procedure related SAEs and non-SAEs by implanter experience Effectiveness · Change from baseline to 1, 3, and 6 months post-permanent implant for overall pain intensity measured using the Visual Analog Scale (VAS) · Change from baseline to 1, 3, and 6 months post-permanent implant for physical function measured using the PROMIS-29 Profile · Change from baseline to 1, 3, and 6 months post-permanent implant for quality of life measured using the PROMIS Global Health Scale · Change from baseline to 1, 3, 6 and 12 months post-permanent implant for neuropathic pain measured using the Neuropathic Pain Scale (NPS) · Change from baseline to 1, 3, 6 and 12 months post-permanent implant for sleep disturbance, anxiety, depression, fatigue, ability to participate in social roles and activities, pain interference and pain intensity measured using the PROMIS-29 Profile · Patient Global Impression of Change (PGIC) at 1, 3, 6 and 12 months post-permanent implant |
| Design |  | This study is a prospective, single arm, observational post-approval study. Study follow-ups will occur at 1, 3, 6, and 12 months post-permanent implant. |
| Devices used | DRG Neurostimulation System with Axium or Proclaim DRG Implantable Pulse Generator |
| Study Population | A maximum of 426 adult subjects with moderate to severe chronic, intractable pain of the lower limbs due to CRPS types I and II will undergo a trial of the neurostimulator system at up to 45 study sites in the United States. |
| Inclusion/Exclusion Criteria | Inclusion Criteria · Subject is male or female > 22 and < 75 years of age. · Subject has moderate to severe chronic intractable pain of the lower limbs resulting from Complex Regional Pain Syndrome (CRPS) types I or II. · Subject has a baseline VAS score of > 60 mm for overall pain at the time of the baseline assessment. · Subject is willing and able to comply with the study requirements. · Subject can provide written informed consent. Exclusion Criteria · Subject has an active implantable medical device including but not limited to cardiac pacemakers and cardiac defibrillators. · Subject is currently involved in medically related litigation, including workers compensation. · Subject has a life expectancy of less than one year. · Subject is pregnant or of childbearing potential and not using adequate contraception as determined by the investigator. · Subject has, or plans to have, a spinal cord stimulation system or infusion pump system implanted. · Subject has, or plans to have, a peripheral nerve stimulation system (PNS) or peripheral nerve field stimulation system (PNfS) implanted. · Subject is considered to be a poor surgical or study candidate, which may include, but is not limited to the following: any medical, social, or psychological problem that could complicate the implant procedure and/or recovery from the implant procedure or could complicate the required procedures and evaluations of the study in the judgment of the investigator. |
| Data Collection | The following activities will occur during the study: Baseline Visit · Subject signs informed consent. · Subject evaluated for eligibility . · Brief treatment history · Pain mapping · Subject completes VAS, NPS, PROMIS-29 Profile, and PROMIS Global Health · Schedule system trial implant Trial System Implant (approximately 30 days from Baseline visit) ·Trial system implantation · Program device · Provide subject with the trial neurostimulator. · Record AEs (if necessary) · Schedule End of System Trial visit End of System Trial Visit (length determined by the site’s standard of care; typically, 3-12 days after trial system implantation) · Document subject trial system evaluation results · Collect programming parameters. · Remove trial leads. · Record AEs (if necessary) · Schedule permanent system implant (if applicable) · Schedule 1-Month safety follow-up call (if necessary) Permanent System Implant (approximately 30 days from End of System Trial visit) · Permanent system implantation · Record AEs (if necessary) · Schedule device activation visit Device Activation (timing determined by the site’s standard of care; may occur at the time of permanent system implant) · Program device · Provide subject with the patient programmer. · Record AEs (if necessary) · Schedule 1-Month follow-up visit 1 Month (± 14 days) Follow-up Visit · Subject completes VAS, NPS, PGIC, PROMIS-29 Profile, and PROMIS Global Health · Stimulation assessment · Collect programming parameters. · Modify programming parameters (if necessary) · Record AEs (if necessary) · Schedule 3-Month follow-up visit 3 Month (± 14 days) Follow-up Visit · Subject completes VAS, NPS, PGIC, PROMIS-29 Profile, and PROMIS Global Health · Stimulation assessment · Collect programming parameters. · Modify programming parameters (if necessary) · Record AEs (if necessary) · Schedule 6-Month follow-up visit 6 Month (± 14 days) Follow-up Visit · Subject completes VAS, NPS, PGIC, PROMIS-29 Profile, and PROMIS Global Health · Stimulation assessment · Collect programming parameters. · Modify programming parameters (if necessary) · Record AEs (if necessary) · Schedule 12-Month follow-up visit 12 Month (± 30 days) Follow-up Visit · Subject completes VAS, NPS, PGIC, PROMIS-29 Profile, and PROMIS Global Health · Stimulation assessment · Collect programming parameters . · Record AEs (if necessary) · Modify programming parameters (if necessary) · Exit subject from study. Revisions, Replacements or Explants (additional surgery) · Revision, replacement, or explant procedure · Modify programming parameters (if necessary) · Record AEs · Schedule 1-Month safety follow-up call (if necessary) Unscheduled Visits (occurring after the permanent implant) · Modify programming parameters (if necessary) · Record AEs (if necessary) |

.

### Basis/Need

### Description of the Activity

The primary objective of this study is to demonstrate that the proportion of serious adverse effects (SAEs) at 12 months for subjects who receive a permanently implanted DRG stimulation system is lower than a pre-specified performance goal. A detailed PMCF Study is described in the Clinical Investigation Plan (SJM-CIP-10113) available in Appendix 1.

### C.1.2.2 Identity of the device including any model number/version

All study devices are in commercial distribution in the United States as aids in the management of CRPS I and II. St. Jude Medical (SJM) received FDA approval for the Axium Neurostimulator System under PMA P150004 on February 11, 2016 and approval for the Proclaim DRG IPG under PMA supplement P150004/S002 on November 28, 2016. The components of the DRG Neurostimulation Systems to be used in this study include the following devices:

Table 1: Description of Devices

| Device Component | Model |
| --- | --- |
| 50 cm SlimTip Trial Lead Kit | MN10350-50A |
| 90 cm SlimTip Trial Lead Kit | MN10350-90A |
| 50 cm SlimTip Implant Lead Kit | MN10450-50A |
| 90 cm SlimTip Implant Lead Kit | MN10450-90A |
| Lead Accessories Kit | MN12050 |
| 22 cm Small Curve Delivery Sheath Kit | MN12150 |
| 22 cm Big Curve Delivery Sheath Kit | MN13650 |
| Implantable Neurostimulator Kit (Axium Neurostimulator System) | MN10200 |
| Proclaim DRG IPG | 3664 |
| Connector Cable | MN11350 |
| 30 cm Small Tunneling Tool Kit | MN11950 |
| Clinician Programmer | MN10700 |
| Patient Programmer | MN10600-02 |
| St. Jude Medical Clinician Programmer App | 3874 |
| St. Jude Medical Patient Controller App | 3875 |
| Trial Neurostimulator | MN10100 |
| DRG External Pulse Generator | 7599 (Base) 7032 (Header) |
| Auxiliary Magnet Kit | MN23300 |
| Programmer Charger Kit | MN23400 |
| Programmer Carrying Case | MN13500 |
| 50 cm Lead Extension Kit | MN10550-50 |
| Curved Needle | MN14000 |
| Tunneling Tool | MN15000 |

SlimTip Lead Kits (MN10350-50A, MN10350-90A, MN10450-50A, MN10450-90A)

The leads are designed for percutaneous introduction into the body using a 14-gauge needle and a set of custom delivery tools provided in each kit. The lead is designed to provide stimulation to the intended DRG. It has four cylindrical electrodes spaced at 5 mm intervals on the distal end. The proximal end fits into a four-conductor connector on the connector cable into a four-conductor connector on the extension lead, or into the header of the implantable neurostimulator. The straight stylet is pre-inserted into the lead and the lead is pre-loaded into the sheath during manufacturing.

The SlimTip trial lead kit and implant lead kit contains the lead and the individual delivery devices required to place the lead(s). The lead comes in two lengths 50cm and 90cm. The straight stylet is intended to assist in steering and positioning the lead within the epidural space and is inserted into the leads during manufacturing. A curved stylet is also included in the kit as an option. The delivery needle is intended to access the epidural space, providing a conduit for the lead, guidewire, and delivery sheaths. The suture anchors are intended to anchor the leads to the soft tissue or skin surface proximal to the distal contacts. The guidewire is optional and is intended to verify that the needle is in the epidural space after using a loss of resistance technique.

The delivery sheaths are intended to help deliver the lead percutaneously into the epidural space and are available in two lengths: (22cm and 30cm) as well as two curve shapes (small curve and big curve). Lead model MN10450-50A is MR Conditional when used in conjunction with Proclaim DRG IPG.

Lead Accessories Kit (MN12050)

The lead accessories kit consists of additional tools to aid lead delivery. A 50cm lead extension is packaged with a torque wrench to enable the physician to extend the length of the lead or to provide a connection from the implant lead to the connector cable. The lead extension is intended for chronic implantation. A connector cable (MN11350) is provided in a separate package to connect the leads or lead extension to the trial neurostimulator.

Tunneling Tool Kit (MN11900)

A 30cm tunneling tool is available with associated accessories to provide a conduit for the lead or lead extension to pass under the skin away from the midline of the spine. The tunneling tool is packaged with two exchangeable tips: a blunt pencil tip and a sharp trocar tip. A straw slide over the tunneling tool and then the steel handle removed. The straw provides the conduit for tunneling. An implantable neurostimulator sizer is packaged with the tunneling tool to allow the physician to properly size the pocket for the stimulator.

Tunneling Tool (Model MN15000)

The tunneling tool is used to provide a conduit for the lead or lead extension to the IPG or away from the midline of the spine. It is packaged with two exchangeable tips (a blunt pencil tip and a sharp trocar tip), a straw, a hex key, and a torque wrench.

Trial Neurostimulator (MN10100)

The external trial neurostimulator provides energy and controls electrical signals delivered to the leads. It has a belt clip and must be worn over a piece of clothing by the subject.

DRG External Pulse Generator (EPG; 7599 & 7032)

The DRG EPG functions in the same way as the MN101000 but is compatible with the same programming platform as the Proclaim DRG.

Implantable Neurostimulator Kit (MN10200)

The implantable neurostimulator provides energy and controls electrical signals delivered to the leads. It is typically implanted by the physician in the abdomen or buttocks.

Proclaim DRG IPG, MR Conditional (Model 3664)

The Proclaim DRG IPG was developed based on the FDA approved Proclaim SCS and Axium DRG pulse generators. It is a 16-channel, 4-port, non-rechargeable, MR Conditional IPG. The internal electronics, battery and external casing (can) are identical to the Proclaim SCS IPG used for SCS indications (Model 3660 and 3665). The Proclaim DRG uses the Axium header design (adapted for the Proclaim IPG cannister) which accommodates up to four DRG leads. Additionally, the firmware capability for Proclaim DRG was developed to match the capabilities of the Axium pulse generator. A pocket sizer is an accessory packaged with the IPG and is used to assess the implant size for the IPG. The Proclaim DRG pocket sizer is made of the same material used in the FDA approved Proclaim SCS pocket sizer, but with a different design to reflect the Proclaim DRG IPG shape and form.

Patient Programmer (MN10600-02)

The patient programmer is set up under the instruction of the clinician or clinical staff to allow the subject to change the amplitude of the energy delivered by the implantable or trial neurostimulator system. The subject carries the programmer with them during daily use.

The patient programmer is powered by an internal rechargeable battery or it is plugged into a power outlet using the power supply cord provided. The patient programmer will be used by the subject to communicate with their implantable or trial neurostimulator system, to monitor the device, and to adjust the stimulation settings within their Investigator’s prescribed limits. The subject is provided with a carrying case (MN13500) for carrying their patient programmer during their participation in this study.

The patient programmer has the following features:

- Turns OFF all stimulation

- Turns stimulation ON or OFF for each body region to be treated

- Adjusts stimulation amplitude for each lead in use

- Shows the subject’s identification information

- Shows Investigator name and clinic name, along with contact information

The patient programmer contains an internal magnet to initiate communication with the implantable or trial neurostimulator system. The patient programmer is designed to be easy to use and establishes two-way communication with the implantable or trial neurostimulator system to control the stimulation delivered to the subject.

Clinician Programmer (MN10700)

The clinician programmer is used by trained staff under the guidance of the clinician to communicate wirelessly with the implantable and/or trial neurostimulator system.

The clinician programmer has the following features:

- Turns OFF all stimulation

- Turns stimulation ON for up to four leads and measures lead impedance

- Allows the user to assign body regions that each individual lead is intended to stimulate, modify electrode configurations and adjust stimulation settings for each lead

- Creates multiple groups of stimulation settings for the subject

- Enters subject and lead identification information, Investigator and clinic name and contact information, and Investigator’s notes

- Performs real time test stimulation to assess the subject’s response for each lead

- Enables subject-controlled therapy and configures subject-controlled therapy settings for each lead

- Acquires identification, diagnostic, and historic information about the neurostimulator device

- Programs configured therapy settings for continuous stimulation

The clinician programmer contains an internal magnet to initiate communication with the implantable or trial neurostimulator device.

Clinician Programmer App (Model 3874 Version 3.3)

The stimulation parameter ranges allowed by Clinician Programmer App (Model 3874 Version 3.3) are the same as the approved Axium system. The IPG stimulation settings are established via the Clinician Programmer (CP) App, which is installed on an off-the-shelf Apple iPad Mini. Different stimulation programs may be stored and selected via the Patient Controller (PC) App, which is installed on an off-the-shelf Apple iPod Touch. The clinician programmer communicates with the pulse generator via Bluetooth.

Patient Controller App (Model 3875 Version 3.3)

The stimulation parameter ranges allowed by the patient controller App (Model 3875 Version 3.3) are the same as the approved Axium system. The IPG stimulation settings are established via the Clinician Programmer (CP) App, which is installed on an off-the-shelf Apple iPad Mini. Different stimulation programs may be stored and selected via the Patient Controller (PC) App, which is installed on an off-the-shelf Apple iPod Touch. The patient controller communicates with the pulse generator via Bluetooth.

### C.1.2.3. Intended use of the device in the investigation

The DRG Neurostimulation System is designed to deliver electrical impulses to the dorsal root ganglia. The system consists of an IPG and leads, accessories, programmers, and extensions that are compatible with the system. The intended medical indication is chronic intractable pain.

### C.1.2.4. Objectives of the study

Primary Objective

The primary objective of this study is to demonstrate that the proportion of serious adverse events (SAEs) at 12 months for subjects who receive the permanent DRG IPG is lower than a pre-specified performance goal.

Secondary Objectives

Secondary objectives of this study are as follows:

- To evaluate change in overall pain intensity

- To evaluate change in physical function

- To evaluate change in quality of life

### C.1.2.5 Study design

This study is a prospective, multicenter, single arm, observational post-approval study with follow-ups at 1, 3, 6 and 12 months post-permanent implant. The clinical study will be conducted at up to 45 centers in the United States. A maximum of 426 subjects will be enrolled to achieve 319 subjects with a permanently implanted system, which will account for expected attrition, to yield 287 subjects for the primary analysis at 12 months post-implant (See Section 10.9 Sample Size for full details). After the last subject with a permanently implanted device reaches the 12-month study visit or crosses the 12-month study visit window, the endpoint analysis will be conducted and reported to the FDA.

After the subject signs the informed consent, s/he will be screened according to the inclusion/exclusion criteria and will undergo a baseline evaluation. Those subjects who meet the criteria for participation will undergo a trial of the DRG Neurostimulation System to assess response to DRG stimulation. Only subjects who report a 50% or greater reduction in overall pain intensity through direct patient-reported percentage of pain relief will receive the permanent implant. Device activation will occur according to the site’s standard of care and may occur at the time of permanent system implant. Subjects will then return to the office for follow-up at 1, 3, 6 and 12 months post-permanent implant. Figure 1 below outlines subject flow through the study.

Figure 1: Subject flow through TARGET study.

### C.1.2.6. Primary and secondary endpoint(s)

The primary endpoint is the 12-month SAE rate for subjects receiving the permanent DRG IPG.

Secondary endpoints include:

- Percent change from baseline to 12 months post-permanent implant for overall pain intensity measured using the Visual Analog Scale (VAS)

- Change from baseline to 12 months post-permanent implant for physical function measured using the PROMIS-29 Profile

- Change from baseline to 12 months post-permanent implant for quality of life measured using the PROMIS Global Health Scale

Descriptive Endpoints

Safety:

- Proportion of subjects with device, procedure or stimulation-related SAEs and non-SAEs among subjects who receive the permanent implant

- Proportion of subjects who have AEs related to the trial implant procedure among subjects who fail the trial neurostimulator phase

- Summary of procedure-related SAEs and non-SAEs by implanter experience

Effectiveness

- Percent change from baseline to 1, 3, and 6 months post-permanent implant for overall pain intensity measured using the Visual Analog Scale (VAS)

- Change from baseline to 1, 3, and 6 months post-permanent implant for physical function measured using the PROMIS-29 Profile

- Change from baseline to 1, 3, and 6 months post-permanent implant for quality of life measured using the PROMIS Global Health Scale

- Change from baseline to 1, 3, 6 and 12 months post-permanent implant for neuropathic pain measured using the Neuropathic Pain Scale (NPS)

- Change from baseline to 1, 3, 6 and 12 months post-permanent implant for sleep disturbance anxiety, depression, fatigue, ability to participate in social roles and activities, pain interference and pain intensity measured using the PROMIS-29 Profile

- Patient Global Impression of Change (PGIC) at 1, 3, 6 and 12 months post-permanent implant

### C.1.2.7. Inclusion/exclusion criteria for subject selection

Inclusion Criteria

To participate in this clinical study, the subject must meet all of the following inclusion criteria:

- Subject is male or female ≥ 22 and ≤ 75 years of age.

- Subject has moderate to severe chronic intractable pain of the lower limbs resulting from Complex Regional Pain Syndrome (CRPS) types I or II.

- Subject has a baseline VAS score of > 60 mm for overall pain at the time of the baseline assessment.

- Subject is willing and able to comply with the study requirements.

- Subject is able to provide written informed consent.

Exclusion Criteria

Subjects are not eligible for this clinical study if they meet any of the following exclusion criteria:

- Subject has an active implantable medical device including but not limited to cardiac pacemakers and cardiac defibrillators.

- Subject is currently involved in medically related litigation, including workers compensation.

- Subject has a life expectancy of less than one year.

- Subject is pregnant or of childbearing potential and not using adequate contraception as determined by the investigator.

- Subject has, or plans to have, a spinal cord stimulation system or infusion pump system implanted.

- Subject has, or plans to have, a peripheral nerve stimulation system (PNS) or peripheral nerve field stimulation system (PNfS) implanted.

- Subject is considered a poor surgical or study candidate, which may include, but is not limited to the following:  any medical, social, or psychological problem that could complicate the implant procedure and/or recovery from the implant procedure or could complicate the required procedures and evaluations of the study in the judgment of the investigator.

### Aim

The aim of the study is to confirm safety and performance of the device post market approval. The primary objective of this study is to demonstrate that the proportion of Serious Adverse Effects (SAEs) at 12 months for subjects who receive a permanently implanted DRG stimulation system is lower than a pre-specified performance goal. The primary endpoint is the 12-month SAE rate for subjects receiving an Axium or Proclaim DRG IPG permanent implant. The performance goal is based on the 12-month SAE rate for Axium subjects in the ACCURATE IDE study, where the observed rate was 10.5%. A non-inferiority margin of 5% is assumed.

### Rationale

The sample size required to reject the null hypothesis for the primary endpoint at the 5% significance level with 80% power assuming a 12-month SAE rate of 10.5% is 287 subjects receiving the permanent Axium or Proclaim DRG IPG. Assuming an attrition rate of 25% from trial to permanent Axium or Proclaim DRG IPG implant, and an attrition rate of 10% from permanent implant to 12-month follow-up, a total of 426 (=287/0.9/0.75) subjects will be implanted with a trial system. The target number of subjects for permanent Axium or Proclaim DRG IPG implants is 319, and the target number of subjects for the primary endpoint analysis is 287.

- .

### Known Limitations

Patients contraindicated for the Axium Neurostimulator System are those who:

- Have an active implantable medical device including but not limited to cardiac pacemakers and cardiac defibrillators.

- Are unable to operate the system.

- Are poor surgical risks.

- Are pregnant.

- Are under the age of 18.

Describe the known limitations of the planned activities such as for example incomplete follow up, missing data, etc.

### Timeline

A timeline for this study is provided below. This timeline may change the actual enrollment rate.

| Time = 0; Month 1: Jun ‘16 FDA Approval | Time = Month 2: Jul ‘16 | Time = Month 3: Aug ‘16 1st enrollment | Time = Month 4: Sept ‘16 | Time = Month 5: Oct ‘16 | Time = Month 6: Nov ‘16 | Time = Month 7: Dec ‘16 | Time = Month 8: Jan ‘17 | Time = Month 9: Feb ‘17 | Time = Month 10: Mar ‘17 | Time = Month 11: Apr ‘17 | Time = Month 12: May ‘17 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Study Initiation |  | 9 sites with IRB approval (Actual) | 9 sites with IRB approval (Actual) | 14 sites with IRB approval (Actual) | 16 sites with IRB approval (Actual) | 20 sites with IRB approval (Actual) | 21 sites with IRB approval (Actual) | 25 sites with IRB approval (Actual) | 26 sites with IRB approval (Actual) | 30 with IRB approval (Actual) | 34 with IRB approval (Actual) |
|  |  | Actual Enrollment=1 | Actual Enrollment=2 | Actual Enrollment=0 | Actual Enrollment=8 | Actual Enrollment=5 | Actual Enrollment=2 | Actual Enrollment=5 | Actual Enrollment=7 | Actual Enrollment=10 | Actual Enrollment=10 |
| Cumulative enrollment |  | 1 | 3 | 3 | 11 | 16 | 18 | 23 | 30 | 40 | 50 |

| Time = Month 13: Jun ‘17 | Time = Month 14: Jul ‘17 | Time = Month 15: Aug ‘17 | Time = Month 16: Sept ‘17 | Time = Month 17: Oct ‘17 | Time = Month 18: Nov ‘17 | Time = Month 19: Dec ‘17 | Time = Month 20: Jan ‘18 | Time = Month 21: Feb ‘18 | Time = Month 22: Mar ‘18 | Time = Month 23: Apr ‘18 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 35 with IRB approval (Actual) | 35 with IRB approval (Actual) | 37 with IRB approval (Actual) | 39 with IRB approval (Actual) | 40 with IRB approval (Actual) |  |  |  |  |  | Up to 41 with IRB Approval |
| Actual Enrollment=7 | Actual Enrollment=10 | Actual Enrollment=16 | Actual Enrollment=13 | Actual Enrollment=19 | Actual Enrollment=13 | Actual Enrollment=8 | Actual Enrollment=6 | Actual Enrollment=12 | Actual Enrollment =13 | Actual Enrollment=17 |
| 57 | 67 | 83 | 96 | 115 | 128 | 136 | 142 | 154 | 167 | 184 |

| Time = Month 24: May ‘18 | Time = Month 25: Jun ‘18 | Time = Month 26: Jul ‘18 | Time = Month 27: Aug ‘18 | Time = Month 28: Sept ‘18 | Time = Month 29: Oct ‘18 | Time = Month 30: Nov ‘18 | Time = Month 31: Dec ‘18 | Time = Month 32: Jan ‘19 | Time = Month 33: Feb ‘19 | Time = Month 34: Mar ‘19 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Up to 43 with IRB Approval | Up to 45 with IRB approval (Actual) |  |  |  |  |  |  |  |  |  |
| Actual Enrollment=13 | Actual Enrollment=7 | Actual Enrollment=13 | Actual Enrollment=15 | Actual Enrollment=13 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 |
| 197 | 204 | 217 | 232 | 245 | 257 | 269 | 281 | 293 | 305 | 317 |

| Time = Month 35: Apr ‘19 | Time = Month 36: May ‘19 | Time = Month 37: Jun ‘19 | Time = Month 38: Jul ‘19 | Time = Month 39: Aug ‘19 | Time = Month 40: Sept ‘19 | Time = Month 41: Oct ‘19 | Time = Month 42: Nov ‘19 | Time = Month 43 Dec ‘19 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Up to 41 with IRB Approval | Up to 43 with IRB Approval | Up to 45 with IRB Approval |  |  |  |  |  |  |
| Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=13 |
| 329 | 341 | 353 | 365 | 377 | 389 | 401 | 413 | 426 |

| Time = Month 53: Oct ‘20 | Time = Month 65: Oct ‘21 | Time = Month 69: Feb ‘22 | Time = Month 70 Mar ‘22 |
| --- | --- | --- | --- |
| All subjects implanted with permanent device | 12-month Follow-ups Complete | Primary Analysis Report | Final Report Study Closed |

Estimated time needed to enroll this subject population.

Site activations are projected to begin three months after protocol approval and continue for 16 months. It is expected that the first subject will be consented three months after the approved protocol is sent to the sites and the first permanent implant will occur approximately two months after the first site is activated. All subjects are expected to be enrolled within 43 months. The total duration of the study is expected to be 70 months.

## Registry

Active surveillance includes activities undertaken by St. Jude Medical to proactively gain insight and data into the actual safety and performance of the device, independent of whether any clinical events have occurred.

| Reference | Title | Date | Version |
| --- | --- | --- | --- |
| Common Specifications |  |  |  |
| Harmonized Standards | ISO 14155:2011(EN) Clinical investigation of medical devices for human subjects – Good clinical practice | 2011 |  |
|  | ISO 14971:2012 Medical devices—Application of risk management to medical devices | 2012 |  |
|  | ISO 13485:2016 Medical devices—Quality management systems—Requirements for regulatory purposes | 2016 |  |
| Guidance Documents | 2020-7 Post-market clinical follow-up (PMCF) Plan Template A guide for manufacturers and notified bodies | April 2020 | - |
|  | 2020-8 Post-market clinical follow-up (PMCF) Evaluation Report Template. A Guide for manufacturers and notified bodies | April 2020 | - |
|  | MEDDEV 2.12/2 rev2 Post market clinical follow-up studies. A guide for manufacturers and notified bodies | January 2012 | 2 |

# Estimated Date of the PMCF Evaluation Report

State when the manufacturer plans to have the first report.

(The timelines shall be defined quarterly or at least yearly.  Coordination of this timing should include alignment from Clinical Affairs on the timing of their PMCF Study Progress Reports.)

APPENDICES

<Remove Appendices if not needed>

Appendix 1.  Abbreviations and Acronyms

| Term | Definition |
| --- | --- |