Post-Market Clinical Follow-up Plan

Axium Neurostimulator System

Approval:

| Role/Function | Printed Name | Signature and Date |
| --- | --- | --- |
| Global Regulatory Operations-Clinical Evaluation (GRO-CE) | Sangeeta Mantripragada | See electronic signature in Windchill |
| Clinical Affairs | TBD | See electronic signature in Windchill |
| Clinical Engineering | Jenner Joseph | See electronic signature in Windchill |
| Regulatory Affairs | Jordan Hanson | See electronic signature in Windchill |
| Medical Affairs | Allen Burton, MD | See electronic signature in Windchill |

# Documentation Information

| Post-Market Clinical Follow-up (PMCF) Plan |  |  |  |  |
| --- | --- | --- | --- | --- |
| PMCF Plan Number: |  | 90330520 |  |  |
| PMCF Plan Date: |  | TBD |  |  |
| PMCF Plan Version: |  | 1.0 |  |  |
| Revision History |  |  |  |  |
| Rev. | Revision Date |  | Description of Change | Revised By |
| A | XX |  | Initial Release | Thomas M. Rasmussen |
| B | XX |  | Change to new PMCF Plan template (88778 Rev A) as per MDCG 2020-7 guidelines | Sangeeta Mantripragada |

Table of Contents

A.	Manufacturer Contact Details	6

B.	Device Description and Specification	7

C.	Activities Related to PMCF:  General and Specific Methods and Procedures	14

C.1	Post-Market Study - TARGET	16

C.1.1.	Basis/Need	21

C.1.2.	Description of the Activity	22

C.1.2.1. Identity of the investigation/study	22

C.1.2.2. Identity of the device including any model number/version	23

C.1.2.4. Objectives of the study	29

C.1.2.5 Study design	30

C.1.2.6. Primary and secondary endpoint(s)	31

C.1.2.7. Inclusion/exclusion criteria for subject selection	33

C.1.2.8. Number of enrolled subjects	34

C.1.2.9. Study population	34

C.1.2.10. Summary of study methods	35

C.1.3.	Aim	36

C.1.4.	Rationale	36

C.1.5.	Known Limitations	41

C.1.6.	Timeline	41

C.2	Post market study – REALITY	43

C.2.1.	Basis / Need	46

C.2.2.	Description of the Activity	46

C.2.2.1. Identity of the investigation / Study	47

C.2.2.2. Identity of the Device including any model number/version	47

C.2.2.3. Intended use of the device in the investigation	48

C.2.2.4 Objectives of the study	48

C.2.2.5. Study design	48

C.2.2.6. Primary and Secondary Endpoints	50

C.2.2.7. Inclusion / Exclusion criteria for subject selection	51

C.2.2.8. Number of enrolled subject selection	52

C.2.2.9. Study population	52

C.2.2.10. Summary of study methods	53

C.2.3.	Aim	53

C.2.4.	Rationale	53

C.2.5.	Known Limitations	55

C.2.6.	Timeline	56

C.3	Registry	56

C.4	Real World Data Analysis	56

C.5	Survey from Health Care Professional	56

C.6	Survey from Patients / Users	56

C.7	Case Reports Review	56

Scientific Literature Screening	56

C.8	56

Other Sources of Clinical Data	57

C.9	57

D.	Evaluation of Clinical Data Relating Equivalent or Similar Devices	57

E.	Reference to the Relevant Parts of the Technical Documentation	58

F.	Reference to Any Applicable Common Specifications, Harmonized Standards or Guidance Documents	59

G.	Estimated Date of the PMCF Evaluation Report	60

Appendix 1.	61

APPENDIX 2.	61

APPENDIX 3.	63

# Tables

Table 1: Description of Devices	23

# Figures

Figure 1: Subject flow through TARGET study.	31

Figure 2: REALITY Study Plan.	50

# Manufacturer Contact Details

| Legal manufacturer name | Abbott Medical |
| --- | --- |
| Address | Abbott Medical 6901 Preston Road Plano, Texas 75024 USA |
| SRN | XX |
| Person responsible for regulatory compliance | XX |
| E-mail | XX |
| Phone | XX |
| Fax | XX |
| Authorized representative (if applicable) | Abbott Medical Coordination Center BVBA The Corporate Village Avenue Da Vincilaan 11, Box F1 Zaventem, B-1935 Belgium |
| Contact person | Dr. Michael Thiel |
| E-mail | Michael.Thiel@abbott.com |
| Phone | + 32 277 46878 |
| Fax | XX |

# Device Description and Specification

| Product or trade name | Axium Neurostimulator System |
| --- | --- |
| Model and type | DRG Neurostimulation Systems |
| General description of the device | The DRG Neurostimulation System is a fully implantable medical device that stimulates neural tissue to produce analgesia for the treatment of chronic intractable pain. The system consists of an IPG and leads, accessories, programmers, and extensions that are compatible with the system. |
| Intended purpose | The DRG Neurostimulation System is designed to deliver electrical impulses to the dorsal root ganglia. The intended medical indication is chronic intractable pain. |
| Intended users | Implanting physicians should be experienced in the diagnosis and treatment of chronic pain syndromes and have undergone surgical and device implantation training. |
| Basic UDI-DI | NA |
| Intended patient population | The intended target population for the system are individuals with chronic, intractable pain or complex regional pain syndrome (Type 1 and Type 2), leading to moderate to severe chronic intractable pain of the lower limbs. |
| Medical condition(s) | Chronic intractable pain |
| Indications | The system is indicated for the management of chronic intractable pain. |
| Contraindications | Patients contraindicated for the Axium Neurostimulator System are those who: - Have an active implantable medical device including but not limited to cardiac pacemakers and cardiac defibrillators - Are unable to operate the system - Are poor surgical risks - Are pregnant Are under the age of 18 |
| Warnings | The following general warnings apply to the Axium Neurostimulator System. • External Defibrillators – Safety for use of external defibrillator discharges on a patient receiving neurostimulation has not been established. • Magnetic Resonance Imaging – The patient should be advised to not undergo any elective magnetic resonance imaging (MRI) with the system in place. Use of MRI in the vicinity of the lead(s) may result in forceful dislodgment of the lead(s), or damage to the Neurostimulator. If a voltage is induced through the lead, it may cause uncomfortable (“jolting” or “shocking”) levels of stimulation or injury to the patient. MRI may cause heating at the lead tip and unintended stimulation could result in tissue damage. • CT Scans – If the patient requires a CT scan, all stimulation should be turned OFF prior to the procedure. If stimulation is not turned off, the patient may experience a momentary increase in stimulation, which may be uncomfortable. After the procedure, stimulation should be turned back on and the system should be checked for proper function. • Ultrasonic Scanning – Ultrasonic equipment may cause mechanical damage to the lead if used directly over the site. • Electrosurgery Devices – Electrosurgery devices should not be used in close proximity to implanted lead(s). Contact between an active lead and the electrosurgical pencil can cause direct stimulation of the contacted nerve and can cause severe injury to the patient. Electrosurgery devices may also damage the lead and cause a loss of stimulation. • Pediatric Use – Use of neurostimulation has not been approved specifically for patients under 18 years of age. • Pregnancy – Safety and effectiveness of neurostimulation for use during pregnancy and nursing have not been established. • Non-Emergency Procedures – The patient must be advised that they must not have non-emergency procedures while they are undergoing trial stimulation. • Emergency Procedures – The patient should be instructed to designate a representative (family member or close friend) to notify any emergency medical personnel of their neurostimulator implant, if emergency care is required. Each patient will be provided with a Medical Alert Card to carry with them that will inform emergency medical personnel of the patient’s implant. The patient should be advised to use caution when undergoing any procedure that could include RF or microwave ablation, defibrillation or cardio version. • Routine Medical Procedures – The patient should be instructed to not undergo dental procedures, diathermy, electrolysis, diagnostic ultrasound, static field therapeutic magnets, diagnostic X-ray, and high output ultrasonic lithotripsy. These procedures may provide interference that can affect TNS or INS device operation or use or damage components of the system that may cause patient harm. If the patient with an INS or TNS device is subsequently given any medical treatment in which an electrical current is passed through his/her body from an external source, either the device should first be deactivated, or care should be taken to monitor the functioning of the neurostimulator during the initial stages of treatment. • Diathermy Therapy – Do not use short-wave diathermy, microwave diathermy, or therapeutic ultrasound diathermy (all now referred to as diathermy) on patients implanted with a neurostimulation system. Energy from diathermy can be transferred through the implanted system and cause tissue damage at the location of the implanted electrodes, resulting in severe injury or death. Diathermy is further prohibited because it may also damage the neurostimulation system components. This damage could result in loss of therapy, requiring additional surgery for system removal and replacement. Injury or damage can occur during diathermy treatment whether the neurostimulation system is turned on or off. All patients are advised to inform their health care professionals that they should not be exposed to diathermy treatment. • Cardiac Pacemakers and Implantable Cardiac Defibrillators (ICD) – It is possible that the device may affect the operation of other implantable devices such as pacemakers or implantable cardiac defibrillators. The physician should be aware of any other implantable devices the patient may have or is scheduled to receive. • Explosive or Flammable Gases – Do not use the patient programmer or clinical programmer to communicate with the INS or TNS in an environment where explosive or flammable gas fumes or vapors are present. The operation of the programmer could cause them to ignite, causing severe burns, injury, or death. • Case Damage – If the INS case is pierced or ruptured, an explosion can occur from the battery chemicals, which can lead to severe burns or even death. • Device Components – The use of non-Spinal Modulation components with the system may result in damage to the system and increased risk of harm to the patient. • Component Disposal – Dispose of leads and neurostimulators per local requirements. Do not crush, puncture, or burn the Neurostimulator because explosion or fire may result. • Exposure to Fluids – Exposure of the external TNS or the Connector Cable to water, body fluids, saline, or cleaning agents can cause corrosion and affect stimulation. If this occurs, dry all components thoroughly prior to lead connection. Do not immerse the external TNS or Connector Cable in fluids. • Device Handling – The patient must be instructed to not remove their Trial Lead(s) or Connector Cable. Manipulation of the components may result in an undesired outcome, such as the patient developing an infection, getting undesirable stimulation, or accidentally turning their stimulation off. The patient must be instructed to not rub or exert pressure on the implantable neurostimulator through the skin as this may cause: lead dislodgement leading to stimulation at the implant site, device inversion leading to the inability to communicate with the device, or skin erosion that can lead to another surgical procedure or possible infection. • Equipment Operation – During stimulation, patients should be advised to not operate potentially dangerous equipment, such as power tools, automobiles, or other motor vehicles. They should not climb ladders or participate in other activities where postural change or an abrupt movement could alter the perception of stimulation intensity and cause patients to fall or lose control of equipment or vehicles or injure others. • Patient Activity – Patients should be advised to limit their activities to low or moderate levels during their trial stimulation period and the first six weeks of implantation of the INS. Failure to do so may result in migration of the leads causing loss of stimulation therapy, muscle stimulation or painful stimulation thereby requiring reoperation to reposition. • Theft Detectors and Metal Screening Devices – Certain types of antitheft devices, such as those used at entrances/exits of department stores, libraries, and other public establishments, and/or airport security screening devices may affect stimulation. It is possible that patients who are implanted with non-adjacent multiple leads and/or patients who are sensitive to low stimulation thresholds may experience a momentary increase in their perceived stimulation, which has been described by some patients as uncomfortable or jolting. It is recommended that patients use caution when approaching such a device and that they request assistance to bypass the device. If they must proceed through the device, patients should turn off the neurostimulator and proceed with caution, being sure to move through the detector quickly. • Restricted Areas – The patient should be warned to seek medical guidance before entering environments which could adversely affect the operation of the implanted device, including areas protected by a warning notice preventing entry by patients fitted with a pacemaker. • Electromagnetic Interference (EMI) – Certain commercial equipment such as resistance welders and induction furnaces, certain communication equipment such as microwave transmitters, linear power amplifiers, and high voltage power lines may generate sufficient EMI to interfere with the INS and TNS System resulting in under stimulation, overstimulation, or difficulty in device communications. Certain medical devices may also cause EMI interference and make it difficult to communicate with the INS or TNS; such devices include bone growth stimulators, transcutaneous electrical nerve stimulation (TENS), and implanted cardiac devices. • Lead Movement – The patient should be instructed to avoid excessive bending, twisting, and stretching, entering hyperbaric chambers above 2.0 atmosphere, and operating the neurostimulator while driving or lifting objects of over 2 kg (5 lbs.) for a minimum of 6 weeks after implantation. These activities may cause lead movement, which can result in under stimulation or overstimulation for the patient. Excessive lead migration may lead to the need for reoperation to replace the leads. • Therapeutic Radiation – Therapeutic radiation may damage the electronic circuitry of an implanted neurostimulation system, although no testing has been performed and no definite information on radiation effects is available. Sources of therapeutic radiation include x-rays, cobalt machines, and linear accelerators. If radiation therapy is required, the area over the implanted INS should be shielded with lead. |
| List and description of any variants and / or configurations covered by this plan | - Trial Neurostimulator models - Implantable Neurostimulator models - Trial Slim Tip Lead, 50cm and 90cm models - Implantable Ball Tip Lead, 50cm and 90cm models - Implantable Slim Tip Lead, 50cm and 90cm models - Clinical Programmer models - Patient Programmer models |
| List of any accessories covered by this plan | - Lead Accessories Kit (MN12050) - Tunneling Tool Kit (MN11900) - qqqqTunneling Tool (Model MN15000) |
| Certificate number | NA |
| CND code(s) | NA |
| Class | AIMD, Class III |
| Classification rule | XX |
| Expected lifetime | XX |
| Novel product | ☐ yes ☒ no |
| Novel related clinical procedure | ☐ yes ☒ no |
| Explanation of any novel features | NA |

# Activities Related to PMCF:  General and Specific Methods and Procedures

A summary table of the different PMCF activities foreseen by the manufacturer is provided below:

| Number | Description | Aim | Rationale and known limitations | Timelines |
| --- | --- | --- | --- | --- |
| C 1 | TARGET: A Post-Approval Study to Evaluate Targeted SCS (DRG) Stimulation for the Management of Moderate to Severe Chronic, Intractable, Pain of the Lower Limbs Due to CRPS Types I and II | The primary objective of this study is to demonstrate that the proportion of serious adverse effects at 12 months for subjects who receive the permanent Implantable Pulse Generator (IPG) is lower than a pre-specified performance goal. | The premarket study (ACCURATE) evaluated spinal column stimulation of the dorsal root ganglion (DRG) with the Axium Neurostimulator System for the treatment of chronic, intractable lower extremity neuropathic pain associated with a diagnosis of Complex Regional Pain Syndrome (CRPS) and Peripheral Causalgia (PC). The ACCURATE study met its primary endpoint and the results of the study provide reasonable assurance of safety and effectiveness of the Axium Neurostimulator System. The Axium Neurostimulator System was approved for sale in the United States by FDA and is indicated for spinal column stimulation via epidural and intra-spinal lead access to the dorsal root ganglion as an aid in the management of moderate to severe chronic intractable pain of the lower limbs in adult patients with CRPS types I and II. The Proclaim Neurostimulation System was subsequently approved for the same indication and is replacing the Axium device in commercial distribution. This post-approval study is a condition of FDA approval. Limitations - TBD | 70 Month |
| C 2 | REALITY: Long-Term Real-World Outcomes Study on Patients Implanted with a Neurostimulator | The objective of the study is to capture long-term, real-world, safety and effectiveness data on Abbott’s market-released neurostimulation systems. | This study is designed to collect long-term follow-up data on the safety and effectiveness of Abbott’s market-released neurostimulation systems in the real-world setting. Limitations: | 13 years |

## Post-Market Study - TARGET

Summary of the study plan

| Title: | TARGET (DRG) Post-Approval Study |
| --- | --- |
| Primary Objective: | The primary objective of this study is to demonstrate that the proportion of serious adverse effects (SAEs) at 12 months for subjects who receive the permanent Implantable Pulse Generator (IPG) is lower than a pre-specified objective performance goal. |
| Secondary Objectives: | - To evaluate change in overall pain intensity - To evaluate change in physical function - To evaluate change in quality of life |
| Primary Endpoint: | The primary endpoint is the 12-month SAE rate for subjects receiving a permanent IPG. |
| Secondary Endpoints: | Percent change from baseline to 12 months post-permanent implant for overall pain intensity measured using the Visual Analog Scale (VAS) Change from baseline to 12 months post-permanent implant for physical function measured using the PROMIS-29 Profile Change from baseline to 12 months post-permanent implant for quality of life measured using the PROMIS Global Health Scale |
| Descriptive Endpoints: | Safety - Proportion of subjects with device, procedure or stimulation-related SAEs and non-SAEs among subjects who receive the permanent implant - Proportion of subjects who have AEs related to the trial implant procedure among subjects who fail the trial neurostimulator phase - Summary of procedure-related SAEs and non-SAEs by implanter experience Effectiveness - Change from baseline to 1, 3, and 6 months post-permanent implant for overall pain intensity measured using the Visual Analog Scale (VAS) - Change from baseline to 1, 3, and 6 months post-permanent implant for physical function measured using the PROMIS-29 Profile - Change from baseline to 1, 3, and 6 months post-permanent implant for quality of life measured using the PROMIS Global Health Scale - Change from baseline to 1, 3, 6 and 12 months post-permanent implant for neuropathic pain measured using the Neuropathic Pain Scale (NPS) - Change from baseline to 1, 3, 6 and 12 months post-permanent implant for sleep disturbance, anxiety, depression, fatigue, ability to participate in social roles and activities, pain interference and pain intensity measured using the PROMIS-29 Profile - Patient Global Impression of Change (PGIC) at 1, 3, 6 and 12 months post-permanent implant |
| Design: | This study is a prospective, single arm, observational post-approval study. Study follow-ups will occur at 1, 3, 6, and 12 months post-permanent implant. |
| Devices used: | DRG Neurostimulation System with Axium or Proclaim DRG Implantable Pulse Generator |
| Study Population | A maximum of 426 adult subjects with moderate to severe chronic, intractable pain of the lower limbs due to CRPS types I and II will undergo a trial of the neurostimulator system at up to 45 study sites in the United States. |
| Inclusion/ Exclusion Criteria | Inclusion Criteria - Subject is male or female > 22 and < 75 years of age. - Subject has moderate to severe chronic intractable pain of the lower limbs resulting from Complex Regional Pain Syndrome (CRPS) types I or II. - Subject has a baseline VAS score of > 60 mm for overall pain at the time of the baseline assessment. - Subject is willing and able to comply with the study requirements. - Subject is able to provide written informed consent. Exclusion Criteria - Subject has an active implantable medical device including but not limited to cardiac pacemakers and cardiac defibrillators. - Subject is currently involved in medically related litigation, including workers compensation. - Subject has a life expectancy of less than one year. - Subject is pregnant or of child-bearing potential and not using adequate contraception as determined by the investigator. - Subject has, or plans to have, a spinal cord stimulation system or infusion pump system implanted. - Subject has, or plans to have, a peripheral nerve stimulation system (PNS) or peripheral nerve field stimulation system (PNfS) implanted. - Subject is considered to be a poor surgical or study candidate, which may include, but is not limited to the following: any medical, social, or psychological problem that could complicate the implant procedure and/or recovery from the implant procedure or could complicate the required procedures and evaluations of the study in the judgment of the investigator. |
| Data Collection | The following activities will occur during the study: Baseline Visit - Subject signs informed consent - Subject evaluated for eligibility - Brief treatment history - Pain mapping - Subject completes VAS, NPS, PROMIS-29 Profile, and PROMIS Global Health - Schedule system trial implant Trial System Implant (approximately 30 days from Baseline visit) - Trial system implantation - Program device - Provide subject with the trial neurostimulator - Record AEs (if necessary) - Schedule End of System Trial visit End of System Trial Visit (length determined by the site’s standard of care; typically 3-12 days after trial system implantation) - Document subject trial system evaluation results - Collect programming parameters - Remove trial leads - Record AEs (if necessary) - Schedule permanent system implant (if applicable) - Schedule 1-Month safety follow-up call (if necessary) Permanent System Implant (approximately 30 days from End of System Trial visit) - Permanent system implantation - Record AEs (if necessary) - Schedule device activation visit Device Activation (timing determined by the site’s standard of care; may occur at the time of permanent system implant) - Program device - Provide subject with the patient programmer - Record AEs (if necessary) - Schedule 1-Month follow-up visit 1 Month (± 14 days) Follow-up Visit - Subject completes VAS, NPS, PGIC, PROMIS-29 Profile, and PROMIS Global Health - Stimulation assessment - Collect programming parameters - Modify programming parameters (if necessary) - Record AEs (if necessary) - Schedule 3-Month follow-up visit 3 Month (± 14 days) Follow-up Visit - Subject completes VAS, NPS, PGIC, PROMIS-29 Profile, and PROMIS Global Health - Stimulation assessment - Collect programming parameters - Modify programming parameters (if necessary) - Record AEs (if necessary) - Schedule 6-Month follow-up visit 6 Month (± 14 days) Follow-up Visit - Subject completes VAS, NPS, PGIC, PROMIS-29 Profile, and PROMIS Global Health - Stimulation assessment - Collect programming parameters - Modify programming parameters (if necessary) - Record AEs (if necessary) - Schedule 12-Month follow-up visit 12 Month (± 30 days) Follow-up Visit - Subject completes VAS, NPS, PGIC, PROMIS-29 Profile, and PROMIS Global Health - Stimulation assessment - Collect programming parameters - Record AEs (if necessary) - Modify programming parameters (if necessary) - Exit subject from study Revisions, Replacements or Explants (additional surgery) - Revision, replacement or explant procedure - Modify programming parameters (if necessary) - Record AEs - Schedule 1-Month safety follow-up call (if necessary) Unscheduled Visits (occurring after the permanent implant) - Modify programming parameters (if necessary) - Record AEs (if necessary) |

### Basis/Need

The safety and efficacy of the use of DRG stimulation for the treatment of patients diagnosed with chronic intractable pain of the lower limbs is supported by data from a prospective, randomized, multicenter, controlled clinical trial. The ACCURATE Investigational Device Exemption (IDE) trial was approved by FDA to assess the safety and efficacy of DRG stimulation with the Axium Neurostimulator System for individuals with chronic lower limb pain associated with Complex Regional Pain Syndrome (CRPS I) or Peripheral Causalgia (CRPS II). In this study, 152 subjects (76 randomized to each the treatment and control group, where treatment was DRG stimulation and control was SCS) were enrolled across 22 investigational sites. The primary endpoint was the percentage of subjects who were considered treatment successes in the treatment group (DRG stimulation) compared to the control group (traditional SCS), where a treatment success was defined as a subject with at least a 50% lower limb(s) pain reduction on a visual analog scale (VAS) compared to baseline in the area of greatest baseline pain during both the trial stimulation phase and at the 3 month follow-up, and the absence of a stimulation-induced neurological deficit. The primary analysis was a test of non-inferiority between the DRG stimulation group compared to the SCS group with a pre-specified 10% non-inferiority margin. The study met the primary endpoint with statistically significantly more treatment successes in the DRG stimulation group compared to the SCS group at 3 months post permanent implant (81.2% vs. 55.7%, respectively; p = 0.0004). This effect persisted at 12 months post implant (74.2% vs. 53.0%; p=0.0047). Secondary and tertiary outcomes including percent change in VAS scores, change in SF-36 scores, change in Profile of Mood States (POMS) scores, and change in Brief Pain Inventory (BPI) scores were also suggestive of improvements in the group who received DRG stimulation compared to the SCS group.

The most commonly reported definitely-related adverse events (AEs) in the group who received DRG stimulation with the Axium Neurostimulator System in the ACCURATE IDE study were pain at the implantable pulse generator (IPG) pocket (14.5% of subjects) and loss of stimulation due to lead migration (11.8% of subjects). There were a total of 21 serious adverse events (SAEs) in 19 subjects reported in this study. The difference in overall rate of SAEs between the Axium and Control groups was not found to be statistically significantly different (Axium 10.5%, Control 14.5%, p=0.6248). Additionally, the overall difference in the rate of SAEs during the system trial (Axium 1.3%, Control 3.9%, p=0.62) and after the permanent implant (Axium 9.2%, Control 10.5%, p=1.0) were not shown to be statistically significantly different. Only 1 serious adverse device effect (SADE; cardiac arrhythmia post-implant) was classified as possibly or definitely related to the device, procedure and/or stimulation in the group, which received the Axium Neurostimulator System. There were no reports of stimulation induced neurological deficits in either the control or Axium group at any point in the study. Collectively, the results from the ACCURATE IDE were determined to support preliminary safety and efficacy of DRG stimulation with the Axium Neurostimulator System and the therapy was approved by the FDA on February 11, 2016. The Proclaim DRG Neurostimulation System was subsequently approved by FDA on November 28, 2016. Approval is contingent upon continued evidence of safety of the therapy.

The current study is a post-approval study designed to demonstrate safety of DRG Neurostimulation Systems with Axium and Proclaim DRG IPG (subsequently referred to as DRG Neurostimulation Systems) and will be conducted in compliance with the Conditions of Approval as agreed upon with the FDA.

### Description of the Activity

The primary objective of this study is to demonstrate that the proportion of serious adverse effects (SAEs) at 12 months for subjects who receive a permanently implanted DRG stimulation system is lower than a pre-specified performance goal. A detailed PMCF Study is described in the Clinical Investigation Plan (SJM-CIP-10113) available in Appendix 1.

### C.1.2.1. Identity of the investigation/study

The results of the clinical study will be submitted for publication, regardless of the outcome. A ‘Publication Agreement’ will be signed between the Principal Investigator and the Sponsor either as a separate Publication Agreement or within the Clinical Trial Agreement. The International Committee of Medical Journal Editors (ICMJE) guidelines on publication will be followed (www.icmje.org). This study will be posted on ClinicalTrials.gov and results will be posted on ClinicalTrials.gov as required.

Up to 45 study sites in the United States will be used for the trial of the neurostimulator system. A list of Clinical Investigational sites and IRB will be kept under a separate cover and is available upon request. It will be part of the CIP.

- Identity of the investigation/study; CIV-ID or single identification number if performed under the Medical device directives or the MDR, then give the CIV ID or single identification number. Add reference detail if the clinical investigation report is available in EUDAMED.

### C.1.2.2. Identity of the device including any model number/version

All study devices are in commercial distribution in the United States as aids in the management of CRPS I and II. St. Jude Medical (SJM) received FDA approval for the Axium Neurostimulator System under PMA P150004 on February 11, 2016 and approval for the Proclaim DRG IPG under PMA supplement P150004/S002 on November 28, 2016. The components of the DRG Neurostimulation Systems to be used in this study include the following devices:

Table 1: Description of Devices

| Device Component | Model |
| --- | --- |
| 50 cm SlimTip Trial Lead Kit | MN10350-50A |
| 90 cm SlimTip Trial Lead Kit | MN10350-90A |
| 50 cm SlimTip Implant Lead Kit | MN10450-50A |
| 90 cm SlimTip Implant Lead Kit | MN10450-90A |
| Lead Accessories Kit | MN12050 |
| 22 cm Small Curve Delivery Sheath Kit | MN12150 |
| 22 cm Big Curve Delivery Sheath Kit | MN13650 |
| Implantable Neurostimulator Kit (Axium Neurostimulator System) | MN10200 |
| Proclaim DRG IPG | 3664 |
| Connector Cable | MN11350 |
| 30 cm Small Tunneling Tool Kit | MN11950 |
| Clinician Programmer | MN10700 |
| Patient Programmer | MN10600-02 |
| St. Jude Medical Clinician Programmer App | 3874 |
| St. Jude Medical Patient Controller App | 3875 |
| Trial Neurostimulator | MN10100 |
| DRG External Pulse Generator | 7599 (Base) 7032 (Header) |
| Auxiliary Magnet Kit | MN23300 |
| Programmer Charger Kit | MN23400 |
| Programmer Carrying Case | MN13500 |
| 50 cm Lead Extension Kit | MN10550-50 |
| Curved Needle | MN14000 |
| Tunneling Tool | MN15000 |

SlimTip Lead Kits (MN10350-50A, MN10350-90A, MN10450-50A, MN10450-90A)

The leads are designed for percutaneous introduction into the body using a 14-gauge needle and a set of custom delivery tools provided in each kit. The lead is designed to provide stimulation to the intended DRG. It has four cylindrical electrodes spaced at 5 mm intervals on the distal end. The proximal end fits into a four-conductor connector on the connector cable into a four-conductor connector on the extension lead, or into the header of the implantable neurostimulator. The straight stylet is pre-inserted into the lead and the lead is pre-loaded into the sheath during manufacturing.

The SlimTip trial lead kit and implant lead kit contains the lead and the individual delivery devices required to place the lead(s). The lead comes in two lengths 50cm and 90cm. The straight stylet is intended to assist in steering and positioning the lead within the epidural space and is inserted into the leads during manufacturing. A curved stylet is also included in the kit as an option. The delivery needle is intended to access the epidural space, providing a conduit for the lead, guidewire, and delivery sheaths. The suture anchors are intended to anchor the leads to the soft tissue or skin surface proximal to the distal contacts. The guidewire is optional and is intended to verify that the needle is in the epidural space after using a loss of resistance technique.

The delivery sheaths are intended to help deliver the lead percutaneously into the epidural space and are available in two lengths: (22cm and 30cm) as well as two curve shapes (small curve and big curve). Lead model MN10450-50A is MR Conditional when used in conjunction with Proclaim DRG IPG.

Lead Accessories Kit (MN12050)

The lead accessories kit consists of additional tools to aid lead delivery. A 50cm lead extension is packaged with a torque wrench to enable the physician to extend the length of the lead or to provide a connection from the implant lead to the connector cable. The lead extension is intended for chronic implantation. A connector cable (MN11350) is provided in a separate package to connect the leads or lead extension to the trial neurostimulator.

Tunneling Tool Kit (MN11900)

A 30cm tunneling tool is available with associated accessories to provide a conduit for the lead or lead extension to pass under the skin away from the midline of the spine. The tunneling tool is packaged with two exchangeable tips: a blunt pencil tip and a sharp trocar tip. A straw slide over the tunneling tool and then the steel handle removed. The straw provides the conduit for tunneling. An implantable neurostimulator sizer is packaged with the tunneling tool to allow the physician to properly size the pocket for the stimulator.

Tunneling Tool (Model MN15000)

The tunneling tool is used to provide a conduit for the lead or lead extension to the IPG or away from the midline of the spine. It is packaged with two exchangeable tips (a blunt pencil tip and a sharp trocar tip), a straw, a hex key, and a torque wrench.

Trial Neurostimulator (MN10100)

The external trial neurostimulator provides energy and controls electrical signals delivered to the leads. It has a belt clip and must be worn over a piece of clothing by the subject.

DRG External Pulse Generator (EPG; 7599 & 7032)

The DRG EPG functions in the same way as the MN101000 but is compatible with the same programming platform as the Proclaim DRG.

Implantable Neurostimulator Kit (MN10200)

The implantable neurostimulator provides energy and controls electrical signals delivered to the leads. It is typically implanted by the physician in the abdomen or buttocks.

Proclaim DRG IPG, MR Conditional (Model 3664)

The Proclaim DRG IPG was developed based on the FDA approved Proclaim SCS and Axium DRG pulse generators. It is a 16-channel, 4-port, non-rechargeable, MR Conditional IPG. The internal electronics, battery and external casing (can) are identical to the Proclaim SCS IPG used for SCS indications (Model 3660 and 3665). The Proclaim DRG uses the Axium header design (adapted for the Proclaim IPG cannister) which accommodates up to four DRG leads. Additionally, the firmware capability for Proclaim DRG was developed to match the capabilities of the Axium pulse generator. A pocket sizer is an accessory packaged with the IPG and is used to assess the implant size for the IPG. The Proclaim DRG pocket sizer is made of the same material used in the FDA approved Proclaim SCS pocket sizer, but with a different design to reflect the Proclaim DRG IPG shape and form.

Patient Programmer (MN10600-02)

The patient programmer is set up under the instruction of the clinician or clinical staff to allow the subject to change the amplitude of the energy delivered by the implantable or trial neurostimulator system. The subject carries the programmer with them during daily use.

The patient programmer is powered by an internal rechargeable battery or it is plugged into a power outlet using the power supply cord provided. The patient programmer will be used by the subject to communicate with their implantable or trial neurostimulator system, to monitor the device, and to adjust the stimulation settings within their Investigator’s prescribed limits. The subject is provided with a carrying case (MN13500) for carrying their patient programmer during their participation in this study.

The patient programmer has the following features:

- Turns OFF all stimulation

- Turns stimulation ON or OFF for each body region to be treated

- Adjusts stimulation amplitude for each lead in use

- Shows the subject’s identification information

- Shows Investigator name and clinic name, along with contact information

The patient programmer contains an internal magnet to initiate communication with the implantable or trial neurostimulator system. The patient programmer is designed to be easy to use and establishes two-way communication with the implantable or trial neurostimulator system to control the stimulation delivered to the subject.

Clinician Programmer (MN10700)

The clinician programmer is used by trained staff under the guidance of the clinician to communicate wirelessly with the implantable and/or trial neurostimulator system.

The clinician programmer has the following features:

- Turns OFF all stimulation

- Turns stimulation ON for up to four leads and measures lead impedance

- Allows the user to assign body regions that each individual lead is intended to stimulate, modify electrode configurations and adjust stimulation settings for each lead

- Creates multiple groups of stimulation settings for the subject

- Enters subject and lead identification information, Investigator and clinic name and contact information, and Investigator’s notes

- Performs real time test stimulation to assess the subject’s response for each lead

- Enables subject-controlled therapy and configures subject-controlled therapy settings for each lead

- Acquires identification, diagnostic, and historic information about the neurostimulator device

- Programs configured therapy settings for continuous stimulation

The clinician programmer contains an internal magnet to initiate communication with the implantable or trial neurostimulator device.

Clinician Programmer App (Model 3874 Version 3.3)

The stimulation parameter ranges allowed by Clinician Programmer App (Model 3874 Version 3.3) are the same as the approved Axium system. The IPG stimulation settings are established via the Clinician Programmer (CP) App, which is installed on an off-the-shelf Apple iPad Mini. Different stimulation programs may be stored and selected via the Patient Controller (PC) App, which is installed on an off-the-shelf Apple iPod Touch. The clinician programmer communicates with the pulse generator via Bluetooth.

Patient Controller App (Model 3875 Version 3.3)

The stimulation parameter ranges allowed by the patient controller App (Model 3875 Version 3.3) are the same as the approved Axium system. The IPG stimulation settings are established via the Clinician Programmer (CP) App, which is installed on an off-the-shelf Apple iPad Mini. Different stimulation programs may be stored and selected via the Patient Controller (PC) App, which is installed on an off-the-shelf Apple iPod Touch. The patient controller communicates with the pulse generator via Bluetooth.

### C.1.2.3. Intended use of the device in the investigation

The DRG Neurostimulation System is designed to deliver electrical impulses to the dorsal root ganglia. The system consists of an IPG and leads, accessories, programmers, and extensions that are compatible with the system. The intended medical indication is chronic intractable pain.

### C.1.2.4. Objectives of the study

Primary Objective

The primary objective of this study is to demonstrate that the proportion of serious adverse events (SAEs) at 12 months for subjects who receive the permanent DRG IPG is lower than a pre-specified performance goal.

Secondary Objectives

Secondary objectives of this study are as follows:

- To evaluate change in overall pain intensity

- To evaluate change in physical function

- To evaluate change in quality of life

### C.1.2.5 Study design

This study is a prospective, multicenter, single arm, observational post-approval study with follow-ups at 1, 3, 6 and 12 months post-permanent implant. The clinical study will be conducted at up to 45 centers in the United States. A maximum of 426 subjects will be enrolled to achieve 319 subjects with a permanently implanted system, which will account for expected attrition, to yield 287 subjects for the primary analysis at 12 months post-implant (See Section 10.9 Sample Size for full details). After the last subject with a permanently implanted device reaches the 12-month study visit or crosses the 12-month study visit window, the endpoint analysis will be conducted and reported to the FDA.

After the subject signs the informed consent, s/he will be screened according to the inclusion/exclusion criteria and will undergo a baseline evaluation. Those subjects who meet the criteria for participation will undergo a trial of the DRG Neurostimulation System to assess response to DRG stimulation. Only subjects who report a 50% or greater reduction in overall pain intensity through direct patient-reported percentage of pain relief will receive the permanent implant. Device activation will occur according to the site’s standard of care and may occur at the time of permanent system implant. Subjects will then return to the office for follow-up at 1, 3, 6 and 12 months post-permanent implant. Figure 1 below outlines subject flow through the study.

Figure 1: Subject flow through TARGET study.

### C.1.2.6. Primary and secondary endpoint(s)

The primary endpoint is the 12-month SAE rate for subjects receiving the permanent DRG IPG.

Secondary endpoints include:

- Percent change from baseline to 12 months post-permanent implant for overall pain intensity measured using the Visual Analog Scale (VAS)

- Change from baseline to 12 months post-permanent implant for physical function measured using the PROMIS-29 Profile

- Change from baseline to 12 months post-permanent implant for quality of life measured using the PROMIS Global Health Scale

Descriptive Endpoints

Safety:

- Proportion of subjects with device, procedure or stimulation-related SAEs and non-SAEs among subjects who receive the permanent implant

- Proportion of subjects who have AEs related to the trial implant procedure among subjects who fail the trial neurostimulator phase

- Summary of procedure-related SAEs and non-SAEs by implanter experience

Effectiveness

- Percent change from baseline to 1, 3, and 6 months post-permanent implant for overall pain intensity measured using the Visual Analog Scale (VAS)

- Change from baseline to 1, 3, and 6 months post-permanent implant for physical function measured using the PROMIS-29 Profile

- Change from baseline to 1, 3, and 6 months post-permanent implant for quality of life measured using the PROMIS Global Health Scale

- Change from baseline to 1, 3, 6 and 12 months post-permanent implant for neuropathic pain measured using the Neuropathic Pain Scale (NPS)

- Change from baseline to 1, 3, 6 and 12 months post-permanent implant for sleep disturbance anxiety, depression, fatigue, ability to participate in social roles and activities, pain interference and pain intensity measured using the PROMIS-29 Profile

- Patient Global Impression of Change (PGIC) at 1, 3, 6 and 12 months post-permanent implant

### C.1.2.7. Inclusion/exclusion criteria for subject selection

Inclusion Criteria

To participate in this clinical study, the subject must meet all of the following inclusion criteria:

- Subject is male or female ≥ 22 and ≤ 75 years of age.

- Subject has moderate to severe chronic intractable pain of the lower limbs resulting from Complex Regional Pain Syndrome (CRPS) types I or II.

- Subject has a baseline VAS score of > 60 mm for overall pain at the time of the baseline assessment.

- Subject is willing and able to comply with the study requirements.

- Subject is able to provide written informed consent.

Exclusion Criteria

Subjects are not eligible for this clinical study if they meet any of the following exclusion criteria:

- Subject has an active implantable medical device including but not limited to cardiac pacemakers and cardiac defibrillators.

- Subject is currently involved in medically-related litigation, including workers compensation.

- Subject has a life expectancy of less than one year.

- Subject is pregnant or of childbearing potential and not using adequate contraception as determined by the investigator.

- Subject has, or plans to have, a spinal cord stimulation system or infusion pump system implanted.

- Subject has, or plans to have, a peripheral nerve stimulation system (PNS) or peripheral nerve field stimulation system (PNfS) implanted.

- Subject is considered a poor surgical or study candidate, which may include, but is not limited to the following:  any medical, social, or psychological problem that could complicate the implant procedure and/or recovery from the implant procedure or could complicate the required procedures and evaluations of the study in the judgment of the investigator.

### C.1.2.8. Number of enrolled subjects

A maximum of 426 adult subjects with moderate to severe chronic, intractable pain of the lower limbs due to CRPS types I and II will undergo a trial of the neurostimulator system at up to 45 study sites in the United States.

### C.1.2.9. Study population

Prior to participating in the clinical study and conducting study-specific procedures, all subjects will be consented, as required by applicable regulations and the center’s IRB. Written informed consent must be obtained from each subject prior to performance of any study related procedures. The informed consent form must be signed and dated by the subject and by the person obtaining the consent. The principal investigator or his/her authorized designee will conduct the Informed Consent Process. This process will include a verbal discussion with the subject on all aspects of the clinical study that are relevant to the subject’s decision to participate in the clinical study. The subject will be provided with the informed consent form that is written in a language that is understandable to the subject. Failure to obtain informed consent from a subject prior to study enrollment should be reported to St. Jude Medical within 5 working days and to the reviewing center’s IRB consistent with the center’s IRB reporting requirements.

Baseline measurements for the subject’s pain symptoms and location will be taken along with a brief treatment history and completion of subject self-reported questionnaires.

The following information will be collected at the baseline visit:

•	Demographics

•	Diagnosis

•	Pain and treatment history

•	VAS score

•	NPS

•	PROMIS-29 Profile domain scores

•	PROMIS Global Health Scale score

•	Protocol Deviations (if applicable)

### C.1.2.10. Summary of study methods

The clinical study will be conducted in accordance with the CIP. All parties participating in the conduct of the clinical study will be qualified by education, training, or experience to perform their tasks. The clinical study will not commence until St. Jude Medical receives written approval from the IRB and relevant regulatory authorities and all required documents have been collected from the site.

All subjects presenting at the investigational site will be screened by a member of the investigational team previously trained on the CIP and delegated to do so.

When the subject’s participation in the clinical study has been completed, the subject will return to medical care as per their physician’s recommendation. The study will be concluded when all sites are closed, and the final report generated by St. Jude Medical has been provided to sites or St. Jude Medical has provided formal documentation of study closure.

### Aim

The aim of the study is to confirm safety and performance of the device post market approval. The primary objective of this study is to demonstrate that the proportion of Serious Adverse Effects (SAEs) at 12 months for subjects who receive a permanently implanted DRG stimulation system is lower than a pre-specified performance goal. The primary endpoint is the 12-month SAE rate for subjects receiving an Axium or Proclaim DRG IPG permanent implant. The performance goal is based on the 12-month SAE rate for Axium subjects in the ACCURATE IDE study, where the observed rate was 10.5%. A non-inferiority margin of 5% is assumed.

### Rationale

Subjects who meet all of the inclusion criteria and none of the exclusion criteria are eligible to participate in this study. All subjects in the clinical study (including system trial failures, those withdrawn from the clinical study and those lost to follow-up) will be accounted for and documented by assigning an identification code linked to their name, alternative identification or contact information. This log will be kept up to date throughout the clinical study by the principal investigator or his/her authorized designee.

The sample size required to reject the null hypothesis for the primary endpoint at the 5% significance level with 80% power assuming a 12-month SAE rate of 10.5% is 287 subjects receiving the permanent Axium or Proclaim DRG IPG. Assuming an attrition rate of 25% from trial to permanent Axium or Proclaim DRG IPG implant, and an attrition rate of 10% from permanent implant to 12-month follow-up, a total of 426 (=287/0.9/0.75) subjects will be implanted with a trial system. The target number of subjects for permanent Axium or Proclaim DRG IPG implants is 319, and the target number of subjects for the primary endpoint analysis is 287.

The primary objective of this study is to demonstrate that the proportion of SAEs for subjects who receive a permanent Axium or Proclaim DRG IPG is lower than a pre-specified performance goal. The primary endpoint is the 12-month SAE rate for subjects receiving an Axium or Proclaim DRG IPG permanent implant. The performance goal is based on the 12-month SAE rate for Axium subjects in the ACCURATE IDE study, where the observed rate was 10.5%. A non-inferiority margin of 5% is assumed. A total of 8 out of 76 (10.5%) subjects experienced an SAE in the ACCURATE IDE study.

A majority (86%; 74/86) of adverse events in the ACCURATE IDE trial occurred within the first 6 months.  Thus, a 12-month follow-up is sufficient to show the long-term safety of DRG stimulation for the management of chronic, intractable, lower extremity pain resulting from CRPS I or II.

Primary and secondary endpoints will be analyzed by comparing subjects who had previously failed SCS for lack of pain relief (i.e. did not receive > 50% pain relief) to those who had not previously failed SCS (i.e. naïve subjects or those who failed SCS due to reasons other than pain relief). Primary, secondary and descriptive endpoints will be analyzed for two distinct cohorts based upon the type of IPG received at permanent implant (Axium or Proclaim).

The primary endpoint is the 12-month SAE rate from permanent implant until 12-month follow-up. The hypothesis is stated as follows:

H0: p ≥ 15.5%

H1: p < 15.5%

where p is the 12-month SAE rate for subjects who receive the permanent Axium or Proclaim DRG IPG (proportion of subjects who experience at least one SAE from permanent implant through 12-month follow-up). The hypothesis will be tested at the 5% significance level. The analysis population will include subjects who received the permanent Axium or Proclaim DRG IPG, who either have an SAE after permanent implant or complete the 12-month follow-up visit.

The primary endpoint analysis will be conducted using Kaplan-Meier method on subjects who received the permanent Axium or Proclaim DRG IPG (primary analysis population). The serious adverse events (SAE) include adverse events which CEC adjudicates as serious events. The 12-month SAE rate will be estimated using Kaplan-Meier method with a 95% upper confidence bound (UCB) estimated using log-log transformation. If the 95% UCB is less than the 15.5%, the null hypothesis will be rejected.

To address the potential impact of the COVID-19 pandemic and to allow for more direct comparison with the performance goal (developed without the consideration of COVID-19), any adverse events that are deemed by the CEC to be COVID-19-related, possibly COVID-19-related, or the result of community-acquired infection will be excluded from the primary endpoint. Subjects who have these events will be censored at the time of the event. A sensitivity analysis will be performed with COVID-19-related or possibly COVID-19-related SAEs and events resulting from community-acquired infection included. Details are provided in the study specific statistical analysis plan.

Secondary Endpoints will be evaluated with respect to following points:

Reduction in pain intensity (VAS)

The endpoint will test whether subjects achieve a clinically relevant change (mean percentage reduction of at least 30%) in VAS score at 12 months from baseline.  Percentage change in VAS score is calculated within each subject as:

% Change in VAS score = (VAS at baseline – VAS at 12-month)/VAS at baseline*100%

The following hypothesis will be tested:

H0: Mean % Change in VAS score ≤ 30%

H1: Mean % Change in VAS score > 30%

The hypothesis will be tested at the 5% significance level. The analysis will be carried out by calculating the 95% lower confidence bound (LCB) on the mean % Change in VAS from baseline to 12-month visit based on a t-distribution.  If the 95% LCB is greater than 30.0%, the null hypothesis is rejected.  The analysis population for this secondary endpoint will include subjects who complete both baseline and 12-month VAS scales, irrespective of IPG types.

Change in Physical Function (PROMIS Physical Function score)

The endpoint will test whether there is a change in mean PROMIS-29 Physical Function domain scores at 12 months from baseline. Change in PROMIS-29 Physical Function domain score is calculated as:

Change in PROMIS-29 Physical Function domain score = (PROMIS-29 Physical Function domain score at 12-Month – PROMIS-29 Physical Function domain score at Baseline)

The following hypothesis will be tested:

H0: Change in PROMIS-29 Physical Function domain score ≤ 0

H1: Change in PROMIS-29 Physical Function domain score > 0

The hypothesis will be tested at the 5% significance level.  The analysis will be carried out by calculating the 95% LCB on the mean change in PROMIS-29 Physical Function domain score from baseline to 12-month visit based on a t-distribution.  If the 95% LCB is greater than 0, the null hypothesis is rejected. The analysis population for this secondary endpoint will include subjects who complete both the baseline and 12-month PROMIS-29 Profile, irrespective of IPG types.

Change in Global Health (PROMIS Global Health score)

The endpoint will test whether there is a change in mean PROMIS Global Health Score from baseline to 12 months. Change in PROMIS Global Health Sore is calculated as:

Change in PROMIS Global Health Score = (PROMIS Global Health Score at 12-Month – PROMIS Global Health Score at Baseline)

The following hypothesis will be tested:

H0: Change in PROMIS Global Health Score ≤ 0

H1: Change in PROMIS Global Health Score > 0

The hypothesis will be tested at the 5% significance level.  The analysis will be carried out by calculating the 95% LCB on the mean change in PROMIS Global Health Score from baseline to 12-month visit based on a t-distribution.  If the 95% LCB is greater than 0, the null hypothesis is rejected. The analysis population for this secondary endpoint will include subjects who complete both baseline and 12-month PROMIS Global Health questionnaire, irrespective of IPG types.

Primary and secondary endpoints will be analyzed by comparing subjects who had previously failed SCS for lack of pain relief (i.e. did not receive > 50% pain relief) to those who had not previously failed SCS (i.e. naïve subjects or those who failed SCS due to reasons other than pain relief).

Primary, secondary and descriptive endpoints will be analyzed for two distinct cohorts based upon the type of IPG received at permanent implant (Axium or Proclaim).

The following descriptive endpoints will be reported:

- Proportion of subjects with device, procedure or stimulation-related SAEs and non-SAEs among subjects who receive the permanent implant

- Proportion of subjects who have AEs related to the trial implant procedure among subjects who fail the trial neurostimulator phase

- Summary of procedure-related SAEs and non-SAEs by implanter experience

- Change from baseline to 1, 3, and 6 months post-permanent implant for overall pain intensity measured using the Visual Analog Scale (VAS)

- Change from baseline to 1, 3, and 6 months post-permanent implant for physical function measured using the PROMIS-29 Profile

- Change from baseline to 1, 3, and 6 months post-permanent implant for quality of life measured using the PROMIS Global Health Scale

- Change from baseline to 1, 3, 6 and 12 months post-permanent implant for neuropathic pain measured using the Neuropathic Pain Scale (NPS)

- Change from baseline to 1, 3, 6 and 12 months post-permanent implant for sleep disturbance, anxiety, depression, fatigue, ability to participate in social roles and activities, pain interference and pain intensity measured using the PROMIS-29 Profile

- Patient Global Impression of Change (PGIC) at 1, 3, 6 and 12 months post-permanent implant

The analysis population for these descriptive endpoints will include subjects with available data.

### Known Limitations

Describe the known limitations of the planned activities such as for example incomplete follow up, missing data, etc.

### Timeline

A timeline for this study is provided below. This timeline may change the actual enrollment rate.

| Time = 0; Month 1: Jun ‘16 FDA Approval | Time = Month 2: Jul ‘16 | Time = Month 3: Aug ‘16 1st enrollment | Time = Month 4: Sept ‘16 | Time = Month 5: Oct ‘16 | Time = Month 6: Nov ‘16 | Time = Month 7: Dec ‘16 | Time = Month 8: Jan ‘17 | Time = Month 9: Feb ‘17 | Time = Month 10: Mar ‘17 | Time = Month 11: Apr ‘17 | Time = Month 12: May ‘17 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Study Initiation |  | 9 sites with IRB approval (Actual) | 9 sites with IRB approval (Actual) | 14 sites with IRB approval (Actual) | 16 sites with IRB approval (Actual) | 20 sites with IRB approval (Actual) | 21 sites with IRB approval (Actual) | 25 sites with IRB approval (Actual) | 26 sites with IRB approval (Actual) | 30 with IRB approval (Actual) | 34 with IRB approval (Actual) |
|  |  | Actual Enrollment=1 | Actual Enrollment=2 | Actual Enrollment=0 | Actual Enrollment=8 | Actual Enrollment=5 | Actual Enrollment=2 | Actual Enrollment=5 | Actual Enrollment=7 | Actual Enrollment=10 | Actual Enrollment=10 |
| Cumulative enrollment |  | 1 | 3 | 3 | 11 | 16 | 18 | 23 | 30 | 40 | 50 |

| Time = Month 13: Jun ‘17 | Time = Month 14: Jul ‘17 | Time = Month 15: Aug ‘17 | Time = Month 16: Sept ‘17 | Time = Month 17: Oct ‘17 | Time = Month 18: Nov ‘17 | Time = Month 19: Dec ‘17 | Time = Month 20: Jan ‘18 | Time = Month 21: Feb ‘18 | Time = Month 22: Mar ‘18 | Time = Month 23: Apr ‘18 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 35 with IRB approval (Actual) | 35 with IRB approval (Actual) | 37 with IRB approval (Actual) | 39 with IRB approval (Actual) | 40 with IRB approval (Actual) |  |  |  |  |  | Up to 41 with IRB Approval |
| Actual Enrollment=7 | Actual Enrollment=10 | Actual Enrollment=16 | Actual Enrollment=13 | Actual Enrollment=19 | Actual Enrollment=13 | Actual Enrollment=8 | Actual Enrollment=6 | Actual Enrollment=12 | Actual Enrollment =13 | Actual Enrollment=17 |
| 57 | 67 | 83 | 96 | 115 | 128 | 136 | 142 | 154 | 167 | 184 |

| Time = Month 24: May ‘18 | Time = Month 25: Jun ‘18 | Time = Month 26: Jul ‘18 | Time = Month 27: Aug ‘18 | Time = Month 28: Sept ‘18 | Time = Month 29: Oct ‘18 | Time = Month 30: Nov ‘18 | Time = Month 31: Dec ‘18 | Time = Month 32: Jan ‘19 | Time = Month 33: Feb ‘19 | Time = Month 34: Mar ‘19 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Up to 43 with IRB Approval | Up to 45 with IRB approval (Actual) |  |  |  |  |  |  |  |  |  |
| Actual Enrollment=13 | Actual Enrollment=7 | Actual Enrollment=13 | Actual Enrollment=15 | Actual Enrollment=13 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 |
| 197 | 204 | 217 | 232 | 245 | 257 | 269 | 281 | 293 | 305 | 317 |

| Time = Month 35: Apr ‘19 | Time = Month 36: May ‘19 | Time = Month 37: Jun ‘19 | Time = Month 38: Jul ‘19 | Time = Month 39: Aug ‘19 | Time = Month 40: Sept ‘19 | Time = Month 41: Oct ‘19 | Time = Month 42: Nov ‘19 | Time = Month 43 Dec ‘19 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Up to 41 with IRB Approval | Up to 43 with IRB Approval | Up to 45 with IRB Approval |  |  |  |  |  |  |
| Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=12 | Estimated Enrollment=13 |
| 329 | 341 | 353 | 365 | 377 | 389 | 401 | 413 | 426 |

| Time = Month 53: Oct ‘20 | Time = Month 65: Oct ‘21 | Time = Month 69: Feb ‘22 | Time = Month 70 Mar ‘22 |
| --- | --- | --- | --- |
| All subjects implanted with permanent device | 12-month Follow-ups Complete | Primary Analysis Report | Final Report Study Closed |

Estimated time needed to enroll this subject population

Site activations are projected to begin three months after protocol approval and continue for 16 months. It is expected that the first subject will be consented three months after the approved protocol is sent to the sites and the first permanent implant will occur approximately two months after the first site is activated. All subjects are expected to be enrolled within 43 months. The total duration of the study is expected to be 70 months.

## Post market study – REALITY

Summary of the study:

| Title: | REALITY: Long-Term Real-World Outcomes Study on Patients Implanted with a Neurostimulator |
| --- | --- |
| Clinical Study Objective: | The objective of the study is to capture long-term, real-world, safety and effectiveness data on Abbott’s market-released neurostimulation systems. |
| Descriptive Endpoints: | - Evaluation of device and procedure related adverse events, deaths, and device deficiencies at each timepoint throughout the study. - Clinical improvement from baseline to each follow up visit assessed by: - PROMIS-29 - Numeric Rating Scale (NRS) for pain intensity - Pain Catastrophizing scale (PCS) - Oswestry Disability Index (ODI) - Pain condition-related medication usage - Transcutaneous Oxygen pressure measurement (TcPO2) - Patient satisfaction - Patient Global Impression of Change (PGIC) for global improvement of the patient. |
| Design: | This is a post-market, prospective, open-label, non-randomized, multi-center study. |
| Devices used: | Devices included in this study are any country-specific market-released Abbott neurostimulation devices, including any future market-approved devices. This study allows for the inclusion of future iterations of Abbott’s neurostimulation systems and expanded indications as they receive approval from the applicable country regulatory authority. |
| Study Population | The study will enroll up to 2,000 subjects from up to 100 participating centers in the U.S, Canada, Europe, Latin America, Australia, and emerging markets. |
| Inclusion/ Exclusion Criteria | Inclusion Criteria - Subject must provide written informed consent prior to any clinical investigation related procedure. - Subject is at least 18 years (or the minimum age required by local law to consent for participation in a clinical investigation) or older at the time of enrollment. - Subject is scheduled to have an Abbott neurostimulation system implanted within 60 days of baseline. - Subject has a baseline (with no stimulation) pain NRS of > 6. Exclusion Criteria - Subject is enrolled, or intends to participate, in a competing clinical study, as determined by Abbott. - Presence of other anatomic or comorbid conditions, or other medical, social, or psychological conditions that, in the investigator’s opinion, could limit the subject’s ability to participate in the clinical investigation or to comply with follow-up requirements. - Subject has or is scheduled to receive an intrathecal pump. - Subject is part of a vulnerable population (section 5.2.2). - Subject has an existing implanted neuromodulation device to address their chronic pain. |
| Data Collection: | The following activities will occur during the study: Baseline/Pre-procedure Baseline visit must be performed within 60 days prior to the implant procedure and after a minimum of 5 consecutive days with stimulation off. Permanent System Implant Implantation of the neurostimulation system will be performed according to the standard operating procedures. Following system implantation, the stimulator will be activated and programmed by trained personnel and/or an Abbott representative either during post-operative recovery or at an office visit in accordance with the physician’s and hospital’s standard operating procedures. Follow-up Assessments Subjects will be followed at 6 months (± 45 days) and at 1-, 2-, 3-, 4-, and 5- years (± 60 days) from the date of implant. Patient Reported Outcome (PRO) Measures The following PRO measures will be collected: - NRS for pain intensity - PROMIS 29 - PCS - ODI (optional) - PGIC Interim Status Check (Telephone Call) Each subject will have a status check via phone call at 9 months (± 30 days) and at 1.5-, 2.5-, 3.5-, and 4.5- years (± 30 days). This status check will be performed to determine current pain level using NRS, report work status, Patient Global Impression of Change and to determine device-related adverse events, adverse events resulting in death, product out of service, and device deficiency if applicable. Unscheduled Visits (occurring after the permanent implant) - Data regarding pain medication usage, NRS, action(s) taken during this visit related to adverse events, device programming should be documented by completing the appropriate CRF as applicable. Any additional CRF, not specified above, may be collected if deemed appropriate by the investigator. - If an ongoing adverse event is reported during the interim status check, then an unscheduled visit should be arranged. - Following an unscheduled visit, the subject should be seen for the next scheduled study visit within window. |

### Basis / Need

Abbott (previously St. Jude Medical) has developed two neurostimulation therapies to address chronic pain: spinal cord stimulation (SCS) and dorsal root ganglion (DRG) stimulation. Abbott’s expanding portfolio of neurostimulation systems includes spinal cord stimulators that can deliver tonic or BurstDRTM stimulation. In the United States (US) and Europe, Abbott’s tonic and BurstDRTM SCS systems are approved and indicated for the management of chronic intractable pain of the trunk and/or limbs.  In the US, Abbott’s DRG stimulation system is approved for the management of chronic intractable pain of the lower limbs in adults with CRPS Type I and Type II. In Europe, these systems are approved for the management of chronic intractable pain.

The European Union Medical Device Regulation (EU-MDR) 2017/745 requires that all device manufacturers provide continuous monitoring of the safety and effectiveness of CE marked devices. Furthermore, companies that previously obtained CE mark for any device or indication by leveraging data from similar products from another manufacturer must now provide their own evidence for the device and specific indication. The study proposed herein is intended to fulfill the requirements set forth in the EU-MDR.

### Description of the Activity

REALITY study is intended to collect short- and long-term safety and effectiveness data on various populations implanted with Abbott’s neurostimulation systems. A detailed PMCF Study is described in the Clinical Investigation Plan (ABT-CIP-10279 Ver. D; Study Name: REALITY) available in Appendix 1.

### C.2.2.1. Identity of the investigation / Study

The data and results from the clinical investigation are the sole property of the Sponsor. The Sponsor shall have the right to access and use all data and results generated during the clinical investigation. The Investigators will not use this clinical investigation-related data without the written consent of the Sponsor for any purpose other than for clinical investigation completion or for generation of publication materials, as referenced in the Clinical Trial Agreement. Single-center results are not allowed to be published or presented before the multi-center results. Any proposals for publications or presentations by the investigators must be reviewed and approved by the Sponsor in a timely manner to enable Sponsor review in compliance with the Sponsor’s publication policy set forth in the Clinical Trial Agreement.

The Sponsor will be responsible for determining whether to register the clinical investigation on www.clinicaltrials.gov or any other clinical trials, in accordance with the International Committee of Medical Journal Editors guidelines, or any other applicable guidelines. In the event Sponsor determines that the clinical investigation should be registered, Sponsor shall be responsible for any such registration and results posting as required by the ClinicalTrials.gov website. Institution and/or Principal Investigator(s) shall not take any action to register the clinical investigation.

Up to 100 sites in US, Canada, Europe, Latin America, Australia, and emerging markets will be used for the study.

- Identity of the investigation/study; CIV-ID or single identification number if performed under the Medical device directives or the MDR, then give the CIV ID or single identification number. Add reference detail if the clinical investigation report is available in EUDAMED.

### C.2.2.2. Identity of the Device including any model number/version

Devices included in this study are any country-specific market-released Abbott neurostimulation devices, including any future market-approved devices. This study allows for the inclusion of future iterations of Abbott’s neurostimulation systems and expanded indications as they receive approval from the applicable country regulatory authority.

Please refer to the country- and device-specific IFU for additional information regarding the devices used in this clinical investigation.

### C.2.2.3. Intended use of the device in the investigation

Indications for use for each neurostimulation device/system is detailed in the country-specific individual Instructions for Use (IFU) document.

### C.2.2.4 Objectives of the study

The objective of the study is to capture long-term, real-world, safety and effectiveness data on Abbott’s market-released neurostimulation systems.

### C.2.2.5. Study design

The REALITY study is a post-market, prospective, non-randomized, multi-center, single-arm, open-label study. This study is designed to collect active post-market surveillance data on Abbott’s market-released neurostimulation systems in accordance with the EU-MDR (2017/745). It is intended to collect short- and long-term safety and effectiveness data on various populations implanted with Abbott’s neurostimulation systems. This study has broad inclusion criteria and minimal exclusion criteria to ensure the results are representative of the real-world use of these devices. Enrollment caps will be implemented to ensure patients from approved indications are represented.

The study will enroll up to 2,000 subjects from up to 100 participating centers in the U.S, Canada, Europe, Latin America, Australia, and emerging markets. Individuals who are scheduled to receive an implantable Abbott neurostimulation system are eligible for study consideration. Subject enrollment is expected to be completed within 7 years; subjects will be followed for 5 years. Study visits will occur at enrollment, baseline, peri-operative period, and at 6 months, 1 year, and every year until the subject has been followed for 5 years. Study phone calls will occur at 9 months and between yearly follow up visits, beginning at 1.5 years. The total duration of the study is expected to be 13 years, including enrollment, data collection from all subjects, and study close out. Figure 2 summarizes the study plan.

The clinical investigation has been designed to involve as little pain, discomfort, fear, and any other foreseeable risk as possible for subjects. [Refer to the Risks Analysis section (Section 7) of the clinical investigation plan for details.]

Figure 2: REALITY Study Plan.

### C.2.2.6. Primary and Secondary Endpoints

The following descriptive endpoints will be reported using summary statistics.

1.	Evaluation of device and procedure related adverse events, deaths, and device deficiencies at each timepoint throughout the study.

2.	Clinical improvement from baseline to each follow up visit assessed by:

a.	PROMIS-29

b.	Numeric Rating Scale (NRS) for pain intensity

c.	Pain Catastrophizing scale (PCS)

d.	Oswestry Disability Index (ODI)*

e.	Pain condition-related medication usage

f.	Transcutaneous Oxygen pressure measurement (TcPO2)*

3.	Patient satisfaction

4.	Patient Global Impression of Change (PGIC) for global improvement of the patient.

Interim and exploratory analyses may be performed on an ad-hoc basis throughout the life of the study.

*ODI and TcPO2 are optional measures for specific indications only.

### C.2.2.7. Inclusion / Exclusion criteria for subject selection

Assessment for general eligibility criteria is based on medical records of the site and interview with a candidate patient. If some of the clinical and laboratory tests are not included in the standard tests, they must be done but only after written informed consent is obtained. Patients must meet ALL the inclusion criteria to be considered for the clinical investigation. If ANY of the exclusion criteria are met, the patient is excluded from the clinical investigation and cannot be enrolled.

Inclusion Criteria

1.	Subject must provide written informed consent prior to any clinical investigation related procedure.

2.	Subject is at least 18 years (or the minimum age required by local law to consent for participation in a clinical investigation) or older at the time of enrollment.

3.	Subject is scheduled to have an Abbott neurostimulation system implanted within 60 days of baseline.

4.	Subject has a baseline (with no stimulation) pain NRS of > 6.

Exclusion Criteria

1.	Subject is enrolled, or intends to participate, in a competing clinical study, as determined by Abbott.

2.	Presence of other anatomic or comorbid conditions, or other medical, social, or psychological conditions that, in the investigator’s opinion, could limit the subject’s ability to participate in the clinical investigation or to comply with follow-up requirements.

3.	Subject has or is scheduled to receive an intrathecal pump.

4.	Subject is part of a vulnerable population (section 5.2.2).

5.	Subject has an existing implanted neuromodulation device to address their chronic pain.

### C.2.2.8. Number of enrolled subject selection

The study will enroll up to 2000 implanted subjects from up to 100 sites in the U.S, Canada, Europe, Latin America, Australia, and emerging markets. Since there are no statistically powered endpoints, there is no required sample size for the study. No site may enroll more than 100 subjects and no more than 30 subjects per chronic pain indication.

### C.2.2.9. Study population

This clinical investigation will enroll adult male and female individuals scheduled to have an Abbott neurostimulation system implanted. Subjects must meet all eligibility criteria and provide written informed consent prior to conducting any investigation-specific procedures not considered standard of care.

Baseline visit must be performed within 60 days prior to the implant procedure and after a minimum of 5 consecutive days with stimulation off.

The following baseline data will be collected:

  - Patient demographics

  - Pain diagnosis (including subject's primary indication for neuromodulation) and pain location

  - Relevant medical history including pain history and prior interventions (e.g. physical therapy, injections, surgeries) to treat the presenting condition

  - Pain condition-related medication usage (e.g. opioids, anti-convulsants, anti-depressants, sleep aids, topicals)

  - Work status and lifestyle information

  - Numeric Rating Scale (NRS) for pain intensity

  - PROMIS-29

  - Pain Catastrophizing Scale (PCS)

  - Transcutaneous Oxygen pressure measurement (TcPO2) (optional)

  - Oswestry Disability Index (ODI) (optional)

  - Procedure- and device-related adverse event (if applicable)

  - Adverse events resulting in death (if applicable)

  - Device deficiency (if applicable)

### C.2.2.10. Summary of study methods

The clinical investigation will be concluded when:

- All sites are closed AND

- The final report has been provided to investigators or the Sponsor has provided formal documentation of clinical investigation closure.

The clinical investigation report will be submitted within one year of the end of the clinical investigation, or if applicable, will be submitted to the appropriate authorities according to current regulatory guidelines.

### Aim

This study is designed to collect long-term follow-up data on the safety and effectiveness of Abbott’s market-released neurostimulation systems in the real-world setting.

### Rationale

This study has broad inclusion criteria and minimal exclusion criteria to ensure the results are representative of the real-world use of these devices. Enrollment caps will be implemented to ensure patients from approved indications are represented. The clinical investigation has been designed to involve as little pain, discomfort, fear, and any other foreseeable risk as possible for subjects.

To avoid bias, investigators are required to screen every patient scheduled to receive an implantable Abbott neurostimulation system. Enrollment of consecutive subjects should be attempted.

Statistical Considerations

Analysis Populations

The analysis will be performed on the Full Analysis Set (FAS), defined as the enrolled subjects who have been implanted with a permanent implant from Abbott’s neurostimulation systems. Deregistered subjects will be excluded from the FAS.

Statistical Analyses - Descriptive Endpoint Analyses

Summary statistics will be provided for all endpoints. Significance values may be provided for comparisons and measures of change as appropriate. Validated questionnaires will be assessed per their respective instructions.

The following descriptive summary statistics for the endpoints may be reported:

1.	Serious adverse device effects (SADEs), adverse device effects (ADEs), deaths, and device deficiencies will be summarized using counts, percentages or Kaplan-Meier survival estimates.

2.	NRS, PCS, PROMIS-29, ODI, TcPO2, and medication usage measured in change from baseline will be summarized using mean, standard deviation, median, minimum, maximum, and a 95% confidence interval.

3.	Patient satisfaction and PGIC at follow up will be summarized using counts and percentages.

The survival time for subjects who experience an adverse event will be calculated as the number of days from implant to the date the adverse event was first observed. Subjects who do not experience adverse event at the time of withdrawal (e.g. due to patient lost to follow-up, patient non-compliance, patient death, all Abbott neurostimulation products explanted/abandoned and patient withdrawn based on family request or investigator request) will be censored at the date of their withdrawal. For subjects that did not experience any adverse event at the time of analysis, the survival time will be censored at the date of analysis.

Standard errors and two-sided 95% confidence intervals for the Kaplan-Meier survival estimates will also be calculated.

Sample Size Calculation and Assumptions

There will be no sample size calculation or power determination.

Timing of Analysis

The final analysis will be conducted, and final report will be written when the last enrolled subjects who either completes the 5-year follow-up visit or the 5-year follow-up visit window has been passed.

Subgroup Analysis

Subgroup analyses may be performed on an ad-hoc basis to understand outcomes for specific indications and patient populations.

Procedures for Accounting for Missing Data

All available safety-related data will be used without adjustment to account for missing data. For effectiveness and other measures, complete case analysis will be used. Sensitivity analysis may be performed on patient reported outcome measures using multiple imputation.

Planned Interim Analysis

Interim analyses may be performed on an ad-hoc basis to assess continuing safety and effectiveness and provide routine updates to applicable authorities as necessary.

Statistical Criteria for Termination

There are no statistical criteria for termination of this clinical investigation.

### Known Limitations

Describe the known limitations of the planned activities such as for example incomplete follow up, missing data, etc.

### Timeline

Subject enrollment is expected to be completed within 7 years; subjects will be followed for 5 years. The total duration of the study is expected to be 13 years, including enrollment, data collection from all subjects, and study close out.

## Registry

Not Applicable

## Real World Data Analysis

Not Applicable

## Survey from Health Care Professional

Not Applicable

## Survey from Patients / Users

Not Applicable

## Case Reports Review

Not Applicable

## Scientific Literature Screening

Not Applicable

## Other Sources of Clinical Data

Not Applicable

# Evaluation of Clinical Data Relating Equivalent or Similar Devices

| Product name of similar device | Senza®, Nevro | RECHARGING RESTORE™, Medtronic | Axium Neurostimulator System |
| --- | --- | --- | --- |
| Intended purpose | The Nevro Spinal Cord Stimulation System is intended to aid in the management of chronic intractable pain of the trunk and/or limbs, including unilateral and bilateral pan associated with the following: 1) Failed Back Surgery Syndrome; 2) Intractable Low Back Pain; 3) Leg Pain | As an aid in the management of chronic intractable pain of the trunk and/or limbs, peripheral vascular disease or refractory angina pectoris | As an aid in the management of chronic intractable pain |
| Intended users | XX | XX | Implanting physicians should be experienced in the diagnosis and treatment of chronic pain syndromes and have undergone surgical and device implantation training. |
| Intended patient population | XX | XX | The Axium Neurostimulator System is intended to treat patients with chronic intractable pain that has been unresponsive to other methods. |
| Medical condition | 1) Failed Back Surgery Syndrome; 2) Intractable Low Back Pain; 3) Leg Pain | chronic intractable pain of the trunk and/or limbs, peripheral vascular disease or refractory angina pectoris | chronic intractable pain |
| Indication | 1) Failed Back Surgery Syndrome; 2) Intractable Low Back Pain; 3) Leg Pain | chronic intractable pain of the trunk and/or limbs, peripheral vascular disease or refractory angina pectoris | chronic intractable pain |
| Reference to clinical data evaluation in the CER | (date, CR002 Ver S, Section 3.2.4 |  |  |

# Reference to the Relevant Parts of the Technical Documentation

There is no relevant information from the clinical evaluation report and/or from the risk management file to be considered in this plan.

1.  Clinical Evaluation Report (CR002_ R and Date>)

Relevant information to be further analyzed and monitored:

☒ No relevant information from the clinical evaluation report to be considered in this plan.

2.  Risk Management File (RM 006, Rev H 2016.09.29)

Relevant information to be further analyzed and monitored:

☒ No relevant information from the risk management file to be considered in this plan.

# Reference to Any Applicable Common Specifications, Harmonized Standards or Guidance Documents

*Verify that the version and date of the reference documents are current and actively applicable.

| Reference | Title | Date | Version |
| --- | --- | --- | --- |
| Common Specifications | XXX Not Applicable |  |  |
| Harmonized Standards | ISO 14155:2011(EN) Clinical investigation of medical devices for human subjects – Good clinical practice | 2011 |  |
|  | ISO 14971:2012 Medical devices—Application of risk management to medical devices | 2012 |  |
|  | ISO 13485:2016 Medical devices—Quality management systems—Requirements for regulatory purposes | 2016 |  |
| Guidance Documents | 2020-7 Post-market clinical follow-up (PMCF) Plan Template A guide for manufacturers and notified bodies | April 2020 | - |
|  | 2020-8 Post-market clinical follow-up (PMCF) Evaluation Report Template. A Guide for manufacturers and notified bodies | April 2020 | - |
|  | MEDDEV 2.12/2 rev2 Post market clinical follow-up studies. A guide for manufacturers and notified bodies | January 2012 | 2 |

# Estimated Date of the PMCF Evaluation Report

The totality of all results of the Post-Market Clinical Follow-up outlined in this plan will be evaluated and documented annually in PMCF Report in line with PMCF Study Progress Reports thereafter. The PMCF Report will be a part the Axium Neurostimulator System Technical Documentation.

The total duration of the study is expected to be 70 months, including enrollment, data collection from all subjects, and study close out.

APPENDICES

# Appendix 1.

Clinical Investigational Plan (CIP)- SJM-CIP-10113; Study name: TARGET

Clinical Investigation Plan -ABT-CIP-10279 Ver. D; Study Name: REALITY

# APPENDIX 2.

CIP Revision History

| Revision History |  |  |  |  |
| --- | --- | --- | --- | --- |
| Amendment Number | Version | Date | Rationale | Details |
| Not Applicable | VA | 04MAR2016 | First release of CIP | NA |
| 1 | VB | 09MAY2016 | Revised to reflect minor changes requested by FDA. | Added abbreviations to Appendix A. |
| 2 | VC | 13JUL2017 | Revised to increase subject enrollment, incorporate a newly released questionnaire, add a newly approved device, and update trial duration to normal standard of care to prevent deviations. | Change in questionnaires, added Proclaim DRG IPG, removed exclusion criterion, change duration of trial period, revised projected study timeline |
| 3 | VC.1 | 2AUG2017 | Administrative changes to clarify inaccuracies, inconsistencies, and reinstate a previously approved descriptive endpoint assessment to maintain data quality. | Edits to resolve inconsistencies in the mention/description of sample size within the document – unified to fit numbers agreed upon with FDA; corrected part number for the tunneling tool; reinstated the NPS (from Rev B) to accommodate the baseline data collected prior to Rev C) |
| 4 | VD | 21MAR 2019 | Revised to reduce sample size, expand primary analysis cohort to include all device types, extend follow-up duration and add new EPG device (FDA correspondence: 21MAY2018) | Changed primary endpoint to device-related serious AEs which allowed reduction in sample size from 376 to 218, adjusted primary analysis cohort to include patients with Axium and/or Proclaim IPG, changed follow-up duration from 12 to 24 months, added DRG EPG; updated safety reporting requirements per new Sponsor policies. |
|  |  |  | Revised to reflect changes requested by FDA in response to Rev D. (FDA correspondence: 10OCT2018) | Changed primary endpoint back to SAEs and sample size justification to previously agreed upon specifications in Rev C; revised sample size to accommodate greater attrition; increased number of investigational sites. |
|  |  |  | Revised to reflect changes requested by FDA in response to Rev D.1 (FDA correspondence: 11JAN2019) | Changed the name of the report to be submitted in 2021 from ‘Primary Analysis Report’ to Final Report’. Additionally, changed the name of the report to be submitted in 2022 from ‘Final Report’ to ‘Long term follow-up report’. |
|  |  |  | Revised to reflect changes requested by FDA in response to Rev D.2 (FDA correspondence: 21MAR 2019) | Included statement that detailed analysis will be performed to compare the demographics and key clinical characteristics of subjects who have withdrawn and those who continue to be enrolled in the study. Clarified that trial failures are screen failures. Changed the name of the report to be submitted in May 2021 from ‘Final Report’ to ‘Primary Analysis Report’. Additionally, changed the name of the report to be submitted in May 2022 from ‘Long-term follow up’ report to Final Report’. |
| 5 | VE | 28OCT2020 | Revised to reflect changes post FDA correspondence 10SEP 2020 | Section 8.10 telemedicine visits allowed for follow-up visits. NRS will be used to collect pain score as VAS must be administered in person. |
|  |  |  | Revised to reflect changes post FDA correspondence 10SEP 2020 | Section 8.13 Validated Questionnaires updated with description of NRS |
|  |  |  | Revised to reflect changes post FDA correspondence 10SEP 2020 | Section 10,1 Primary endpoint. Endpoint analysis changed to use the Kaplan-Meier method. Details added for COVID-19 events. |
|  |  | 02 FEB 2021 | 18M and 24M visits no longer required | Extended visits added previously were removed to reflect changing reimbursement needs. |

# APPENDIX 3.

Abbreviations and Acronyms

| Term | Definition |
| --- | --- |
| AE | Adverse Event |
| CEC | Clinical Events Committee |
| CIP | Clinical Investigational Plan |
| CRPS | Complex Regional Pain Syndrome |
| DMP | Data Management Plan |
| DRG | Dorsal Root Ganglion Stimulation |
| eCRF | Electronic Case Report Form |
| EDC | Electronic Data Capture |
| ICMJE | International Committee of Medical Journal Editors |
| IPG | Implantable Pulse Generator |
| IRB | Institutional Review Board |
| ISO | International Organization for Standardization |
| MP | Monitoring Plan |
| NA | Not Applicable |
| PG | Pulse Generator |
| RDC | Remote Data Capture |
| SAE | Serious Adverse Event |
| SJM | St. Jude Medical |
| TENS | Transcutaneous Electrical Nerve Stimulation |

| Term | Definition |
| --- | --- |